Integrating the Healthcare Enterprise

IHE Radiation Oncology

Technical Framework Supplement

Treatment Planning – Plan Content

(TPPC)

Draft for Trial Implementation

Date:		July 16, 2018 (1.4)

Author:	IHE Radiation Oncology Technical Committee

Email:		ro@ihe.net

Please verify you have the most recent version of this document. See here for Trial Implementation and Final Text versions and here for Public Comment versions.

Foreword

This is a supplement to the IHE Radiation Oncology Technical Framework V1.8. Each supplement undergoes a process of public comment and trial implementation before being incorporated into the volumes of the Technical Frameworks.

This supplement is published on October 23, 2015 for Public Comment. Comments are invited and can be submitted at http://www.ihe.net/Radiation_Oncology_Public_Comments/. In order to be considered in development of the Trial Implementation version of the supplement, comments must be received by November 23, 2015.

This supplement describes changes to the existing technical framework documents.

“Boxed” instructions like the sample below indicate to the Volume Editor how to integrate the relevant section(s) into the relevant Technical Framework volume.

Amend Section X.X by the following:

Where the amendment adds text, make the added text bold underline. Where the amendment removes text, make the removed text bold strikethrough. When entire new sections are added, introduce with editor’s instructions to “add new text” or similar, which for readability are not bolded or underlined.

General information about IHE can be found at: www.ihe.net.

Information about the IHE Radiation Oncology domain can be found at: ihe.net/IHE_Domains.

Information about the organization of IHE Technical Frameworks and Supplements and the process used to create them can be found at: http://ihe.net/IHE_Process and http://ihe.net/Profiles.

The current version of the IHE Radiation Oncology Technical Framework can be found at: http://ihe.net/Technical_Frameworks.

CONTENTS

Introduction to this Supplement	14

Open Issues and Questions	14

Closed Issues	15

Volume 1 – Profiles	18

1.7	History of Annual Changes	18

1.9	Copyright Permission	18

6	Treatment Planning – Plan Content Integration (TPPC) Profile	19

6.1	TPPC Actors, Transactions, and Content Modules	19

6.1.1 Actor Descriptions and Actor Profile Requirements	23

6.2 TPPC Transaction Options	23

6.2.1	Producer / Consumer Transaction Groupings	26

6.3 TPPC Actor Required Groupings	26

6.4 TPPC Document Content Module	27

6.5 TPPC Overview	28

6.5.1 Concepts	28

6.5.2 Use Case #1: Treatment Replanning	30

6.5.2.1 Treatment Re-planning Use Case Description	30

6.5.2.2 Treatment Replanning Process Flow	30

6.5.3 Use Case #2: TMS Upload	30

6.5.3.1 TMS Upload Use Case Description	31

6.5.3.2 TMS Upload Process Flow	31

6.6	TPPC Security Considerations	32

6.7	TPPC Cross Profile Considerations	32

Appendices	33

Appendix A Actor Summary Definitions	33

Transaction Summary Definitions	35

Glossary	39

Volume 2 – Transactions	40

3.19 TPPC-01 Basic Static Beam Storage	40

3.19.1 Scope	40

3.19.2 Use Case Roles	40

3.19.3 Referenced Standards	40

3.19.4 Interaction Diagram	41

3.19.4.1 Basic Static Beam Storage	41

3.19.4.1.1 Trigger Events	41

3.19.4.1.2 Message Semantics	41

3.19.4.1.2.1 Storage of RT Plan containing a Basic Static Beam	41

3.19.4.1.2.2 Optional Modifiers	42

3.19.4.1.3 Expected Actions	42

3.19.5 Security Considerations	42

3.20 TPPC-02: Basic Static Beam Retrieval	42

3.20.1 Scope	42

3.20.2 Use Case Roles	42

3.20.3 Referenced Standards	43

3.20.4 Interaction Diagram	43

3.20.4.1 Basic Static Beam Retrieval	43

3.20.4.1.1 Trigger Events	43

3.20.4.1.2 Message Semantics	43

3.20.4.1.2.1 Storage of RT Plan containing a Basic Static Beam	43

3.20.4.1.2.2 Optional Modifiers	44

3.20.4.1.3 Expected Actions	44

3.20.5 Security Considerations	44

3.21 TPPC-03 Basic Static MLC Beam Storage	44

3.21.1 Scope	44

3.21.2 Use Case Roles	44

3.21.3 Referenced Standards	45

3.21.4 Interaction Diagram	45

3.21.4.1 Basic Static MLC Beam Storage	45

3.21.4.1.1 Trigger Events	45

3.21.4.1.2 Message Semantics	45

3.21.4.1.2.1 Storage of RT Plan containing a Basic Static MLC Beam	45

3.21.4.1.2.2 Optional Modifiers	46

3.21.4.1.3 Expected Actions	46

3.21.5 Security Considerations	46

3.22 TPPC-04: Basic Static MLC Beam Retrieval	46

3.22.1 Scope	46

3.22.2 Use Case Roles	46

3.22.3 Referenced Standards	47

3.22.4 Interaction Diagram	47

3.22.4.1 Basic Static MLC Beam Retrieval	47

3.22.4.1.1 Trigger Events	47

3.22.4.1.2 Message Semantics	47

3.22.4.1.2.1 Storage of RT Plan containing a Basic Static MLC Beam	47

3.22.4.1.2.2 Optional Modifiers	48

3.22.4.1.3 Expected Actions	48

3.22.5 Security Considerations	48

3.23 TPPC-05 Arc Beam Storage	48

3.23.1 Scope	48

3.23.2 Use Case Roles	48

3.23.3 Referenced Standards	49

3.23.4 Interaction Diagram	49

3.23.4.1 Arc Beam Storage	49

3.23.4.1.1 Trigger Events	49

3.23.4.1.2 Message Semantics	49

3.23.4.1.2.1 Storage of RT Plan containing an Arc Beam	49

3.23.4.1.2.2 Optional Modifiers	50

3.23.4.1.3 Expected Actions	50

3.23.5 Security Considerations	50

3.24 TPPC-06: Arc Beam Retrieval	50

3.24.1 Scope	50

3.24.2 Use Case Roles	50

3.24.3 Referenced Standards	51

3.24.4 Interaction Diagram	51

3.24.4.1 Arc Beam Retrieval	51

3.24.4.1.1 Trigger Events	51

3.24.4.1.2 Message Semantics	51

3.24.4.1.2.1 Storage of RT Plan containing an Arc Beam	51

3.24.4.1.2.2 Optional Modifiers	52

3.24.4.1.3 Expected Actions	52

3.24.5 Security Considerations	52

3.25 TPPC-07 MLC Fixed Aperture Arc Beam Storage	52

3.25.1 Scope	52

3.25.2 Use Case Roles	52

3.25.3 Referenced Standards	53

3.25.4 Interaction Diagram	53

3.25.4.1 MLC Fixed Aperture Arc Beam Storage	53

3.25.4.1.1 Trigger Events	53

3.25.4.1.2 Message Semantics	53

3.25.4.1.2.1 Storage of RT Plan containing a MLC Fixed Aperture Arc Beam	53

3.25.4.1.2.2 Optional Modifiers	54

3.25.4.1.3 Expected Actions	54

3.25.5 Security Considerations	54

3.26 TPPC-08: MLC FIXED APERTURE ARC Beam Retrieval	54

3.26.1 Scope	54

3.26.2 Use Case Roles	54

3.26.3 Referenced Standards	55

3.26.4 Interaction Diagram	55

3.26.4.1 MLC Fixed Aperture Arc Beam Retrieval	55

3.26.4.1.1 Trigger Events	55

3.26.4.1.2 Message Semantics	55

3.26.4.1.2.1 Storage of RT Plan containing a MLC Fixed Aperture Arc Beam	55

3.26.4.1.2.2 Optional Modifiers	56

3.26.4.1.3 Expected Actions	56

3.26.5 Security Considerations	56

3.27 TPPC-09 MLC Variable Aperture Arc Beam Storage	56

3.27.1 Scope	56

3.27.2 Use Case Roles	56

3.27.3 Referenced Standards	57

3.27.4 Interaction Diagram	57

3.27.4.1 MLC Variable Aperture Arc Beam Storage	57

3.27.4.1.1 Trigger Events	57

3.27.4.1.2 Message Semantics	57

3.27.4.1.2.1 Storage of RT Plan containing a MLC Variable Aperture Arc Beam	57

3.27.4.1.2.2 Optional Modifiers	58

3.27.4.1.3 Expected Actions	58

3.27.5 Security Considerations	58

3.28 TPPC-10: MLC Variable Aperture Arc Beam Retrieval	58

3.28.1 Scope	58

3.28.2 Use Case Roles	58

3.28.3 Referenced Standards	59

3.28.4 Interaction Diagram	59

3.28.4.1 MLC Variable Aperture Arc Beam Retrieval	59

3.28.4.1.1 Trigger Events	59

3.28.4.1.2 Message Semantics	59

3.28.4.1.2.1 Storage of RT Plan containing a MLC Variable Aperture Arc Beam	59

3.28.4.1.2.2 Optional Modifiers	60

3.28.4.1.3 Expected Actions	60

3.28.5 Security Considerations	60

3.29 TPPC-11 Hard Wedge Beam Storage	60

3.29.1 Scope	60

3.29.2 Use Case Roles	60

3.29.3 Referenced Standards	61

3.29.4 Interaction Diagram	61

3.29.4.1 Hard Wedge Beam Storage	61

3.29.4.1.1 Trigger Events	61

3.29.4.1.2 Message Semantics	61

3.29.4.1.2.1 Storage of RT Plan containing a Hard Wedge Beam	61

3.29.4.1.2.2 Optional Modifiers	62

3.29.4.1.3 Expected Actions	62

3.29.5 Security Considerations	62

3.30 TPPC-12: Hard Wedge Beam Retrieval	62

3.30.1 Scope	62

3.30.2 Use Case Roles	62

3.30.3 Referenced Standards	63

3.30.4 Interaction Diagram	63

3.30.4.1 Hard Wedge Beam Retrieval	63

3.30.4.1.1 Trigger Events	63

3.30.4.1.2 Message Semantics	63

3.30.4.1.2.1 Storage of RT Plan containing a Hard Wedge Beam	63

3.30.4.1.2.2 Optional Modifiers	64

3.30.4.1.3 Expected Actions	64

3.30.5 Security Considerations	64

3.31 TPPC-13 Virtual Wedge Beam Storage	64

3.31.1 Scope	64

3.31.2 Use Case Roles	64

3.31.3 Referenced Standards	65

3.31.4 Interaction Diagram	65

3.31.4.1 Virtual Wedge Beam Storage	65

3.31.4.1.1 Trigger Events	65

3.31.4.1.2 Message Semantics	65

3.31.4.1.2.1 Storage of RT Plan containing a Virtual Wedge Beam	65

3.31.4.1.2.2 Optional Modifiers	66

3.31.4.1.3 Expected Actions	66

3.31.5 Security Considerations	66

3.32 TPPC-14: Virtual Wedge Beam Retrieval	66

3.32.1 Scope	66

3.32.2 Use Case Roles	66

3.32.3 Referenced Standards	67

3.32.4 Interaction Diagram	67

3.32.4.1 Virtual Wedge Beam Retrieval	67

3.32.4.1.1 Trigger Events	67

3.32.4.1.2 Message Semantics	67

3.32.4.1.2.1 Storage of RT Plan containing a Virtual Wedge Beam	67

3.32.4.1.2.2 Optional Modifiers	68

3.32.4.1.3 Expected Actions	68

3.32.5 Security Considerations	68

3.33 TPPC-15 Motorized Wedge Beam Storage	68

3.33.1 Scope	68

3.33.2 Use Case Roles	68

3.33.3 Referenced Standards	69

3.33.4 Interaction Diagram	69

3.33.4.1 Motorized Wedge Beam Storage	69

3.33.4.1.1 Trigger Events	69

3.33.4.1.2 Message Semantics	69

3.33.4.1.2.1 Storage of RT Plan containing a Motorized Wedge Beam	69

3.33.4.1.2.2 Optional Modifiers	70

3.33.4.1.3 Expected Actions	70

3.33.5 Security Considerations	70

3.34 TPPC-16: Motorized Wedge Beam Retrieval	70

3.34.1 Scope	70

3.34.2 Use Case Roles	70

3.34.3 Referenced Standards	71

3.34.4 Interaction Diagram	71

3.34.4.1 Motorized Wedge Beam Retrieval	71

3.34.4.1.1 Trigger Events	71

3.34.4.1.2 Message Semantics	71

3.34.4.1.2.1 Storage of RT Plan containing a Motorized Wedge Beam	71

3.34.4.1.2.2 Optional Modifiers	72

3.34.4.1.3 Expected Actions	72

3.34.5 Security Considerations	72

3.35 TPPC-17 Static Electron Beam Storage	72

3.35.1 Scope	72

3.35.2 Use Case Roles	72

3.35.3 Referenced Standards	73

3.35.4 Interaction Diagram	73

3.35.4.1 Static Electron Beam Storage	73

3.35.4.1.1 Trigger Events	73

3.35.4.1.2 Message Semantics	73

3.35.4.1.2.1 Storage of RT Plan containing a Static Electron Beam	74

3.35.4.1.2.2 Optional Modifiers	74

3.35.4.1.3 Expected Actions	74

3.35.5 Security Considerations	74

3.36 TPPC-18: Static Electron Beam Retrieval	74

3.36.1 Scope	74

3.36.2 Use Case Roles	74

3.36.3 Referenced Standards	75

3.36.4 Interaction Diagram	75

3.36.4.1 Static Electron Beam Retrieval	75

3.36.4.1.1 Trigger Events	75

3.36.4.1.2 Message Semantics	75

3.36.4.1.2.1 Storage of RT Plan containing a Static Electron Beam	75

3.36.4.1.2.2 Optional Modifiers	76

3.36.4.1.3 Expected Actions	76

3.36.5 Security Considerations	76

3.37 TPPC-19 Step & Shoot Beam Storage	76

3.37.1 Scope	76

3.37.2 Use Case Roles	76

3.37.3 Referenced Standards	77

3.37.4 Interaction Diagram	77

3.37.4.1 Step & Shoot Beam Storage	77

3.37.4.1.1 Trigger Events	77

3.37.4.1.2 Message Semantics	77

3.37.4.1.2.1 Storage of RT Plan containing a Step & Shoot Beam	77

3.37.4.1.2.2 Optional Modifiers	78

3.37.4.1.3 Expected Actions	78

3.37.5 Security Considerations	78

3.38 TPPC-20: Step & Shoot Beam Retrieval	78

3.38.1 Scope	78

3.38.2 Use Case Roles	78

3.38.3 Referenced Standards	79

3.38.4 Interaction Diagram	79

3.38.4.1 Step & Shoot Beam Retrieval	79

3.38.4.1.1 Trigger Events	79

3.38.4.1.2 Message Semantics	79

3.38.4.1.2.1 Storage of RT Plan containing a Step & Shoot Beam	79

3.38.4.1.2.2 Optional Modifiers	80

3.38.4.1.3 Expected Actions	80

3.38.5 Security Considerations	80

3.39 TPPC-21 Sliding Window Beam Storage	80

3.39.1 Scope	80

3.39.2 Use Case Roles	80

3.39.3 Referenced Standards	81

3.39.4 Interaction Diagram	81

3.39.4.1 Sliding Window Beam Storage	81

3.39.4.1.1 Trigger Events	81

3.39.4.1.2 Message Semantics	81

3.39.4.1.2.1 Storage of RT Plan containing a Sliding Window Beam	81

3.39.4.1.2.2 Optional Modifiers	82

3.39.4.1.3 Expected Actions	82

3.39.5 Security Considerations	82

3.40 TPPC-22: Sliding Window Beam Retrieval	82

3.40.1 Scope	82

3.40.2 Use Case Roles	82

3.40.3 Referenced Standards	83

3.40.4 Interaction Diagram	83

3.40.4.1 Sliding Window Beam Retrieval	83

3.40.4.1.1 Trigger Events	83

3.40.4.1.2 Message Semantics	83

3.40.4.1.2.1 Storage of RT Plan containing a Sliding Window Beam	83

3.40.4.1.2.2 Optional Modifiers	83

3.40.4.1.3 Expected Actions	84

3.40.5 Security Considerations	84

3.41 TPPC-23 IMAT/VMAT Beam Storage	84

3.41.1 Scope	84

3.41.2 Use Case Roles	84

3.41.3 Referenced Standards	84

3.41.4 Interaction Diagram	85

3.41.4.1 IMAT/VMAT Beam Storage	85

3.41.4.1.1 Trigger Events	85

3.41.4.1.2 Message Semantics	85

3.41.4.1.2.1 Storage of RT Plan containing a IMAT/VMAT Beam	85

3.41.4.1.2.2 Optional Modifiers	86

3.41.4.1.3 Expected Actions	86

3.41.5 Security Considerations	86

3.42 TPPC-24: IMAT/VMAT Beam Retrieval	86

3.42.1 Scope	86

3.42.2 Use Case Roles	86

3.42.3 Referenced Standards	87

3.42.4 Interaction Diagram	87

3.42.4.1 IMAT/VMAT Beam Retrieval	87

3.42.4.1.1 Trigger Events	87

3.42.4.1.2 Message Semantics	87

3.42.4.1.2.1 Storage of RT Plan containing a IMAT/VMAT Beam	87

3.42.4.1.2.2 Optional Modifiers	87

3.42.4.1.3 Expected Actions	88

3.42.5 Security Considerations	88

3.43 TPPC-25 Photon Applicator Beam Storage	88

3.43.1 Scope	88

3.43.2 Use Case Roles	88

3.43.3 Referenced Standards	88

3.43.4 Interaction Diagram	89

3.43.4.1 Photon Applicator Beam Storage	89

3.43.4.1.1 Trigger Events	89

3.43.4.1.2 Message Semantics	89

3.43.4.1.2.1 Storage of RT Plan containing a Photon Applicator Beam	89

3.43.4.1.2.2 Optional Modifiers	90

3.43.4.1.3 Expected Actions	90

3.43.5 Security Considerations	90

3.44 TPPC-26: Photon Applicator Beam Retrieval	90

3.44.1 Scope	90

3.44.2 Use Case Roles	90

3.44.3 Referenced Standards	91

3.44.4 Interaction Diagram	91

3.44.4.1 Photon Applicator Beam Retrieval	91

3.44.4.1.1 Trigger Events	91

3.44.4.1.2 Message Semantics	91

3.44.4.1.2.1 Storage of RT Plan containing a Photon Applicator Beam	91

3.44.4.1.2.2 Optional Modifiers	91

3.44.4.1.3 Expected Actions	92

3.44.5 Security Considerations	92

3.45 TPPC-27 Photon Applicator Arc Beam Storage	92

3.45.1 Scope	92

3.45.2 Use Case Roles	92

3.45.3 Referenced Standards	92

3.45.4 Interaction Diagram	93

3.45.4.1 Photon Applicator Arc Beam Storage	93

3.45.4.1.1 Trigger Events	93

3.45.4.1.2 Message Semantics	93

3.45.4.1.2.1 Storage of RT Plan containing a Photon Applicator Arc Beam	93

3.45.4.1.2.2 Optional Modifiers	94

3.45.4.1.3 Expected Actions	94

3.45.5 Security Considerations	94

3.46 TPPC-28: Photon Applicator Arc Beam Retrieval	94

3.46.1 Scope	94

3.46.2 Use Case Roles	94

3.46.3 Referenced Standards	95

3.46.4 Interaction Diagram	95

3.46.4.1 Photon Applicator Arc Beam Retrieval	95

3.46.4.1.1 Trigger Events	95

3.46.4.1.2 Message Semantics	95

3.46.4.1.2.1 Storage of RT Plan containing a Photon Applicator Beam	95

3.46.4.1.2.2 Optional Modifiers	95

3.46.4.1.3 Expected Actions	96

3.46.5 Security Considerations	96

Volume 3 – Content Modules	97

5 Namespaces and Vocabularies	98

6 Content Modules	99

7	DICOM Content Definition	100

7.1	Conventions	100

7.2	General Definitions	100

7.3	IOD Definitions	100

7.3.1	Prescription IODs	100

7.3.2	Plan IODs	100

7.3.2.1	RT Plan IOD for Photon External Beam in Planning State	100

7.3.2.1.1	Referenced Standards	100

7.3.2.1.2	IOD Definition	100

7.4	Module Definitions	102

7.4.1	General Modules	102

7.4.1.1	Patient Module	102

7.4.1.1.1	Patient Module Base Content	102

7.4.1.2	Study Module	102

7.4.1.2.1	Study Module Base Content	102

7.4.1.3	General Series Module	102

7.4.1.3.1	General Series Module Base Content	102

7.4.1.4	RT Series Module	102

7.4.1.4.1	RT Series Module Base Content	102

7.4.1.5	Equipment Module	102

7.4.1.5.1	Equipment Module Base Content	102

7.4.1.6	SOP Common Module	102

7.4.1.6.1	SOP Common Module Base Content	102

7.4.1.7	Frame of Reference Module	102

7.4.1.7.1	Frame of Reference Module Base Content	102

7.4.2	Workflow-Related Modules	103

7.4.3	General Plan-Related Modules	103

7.4.3.1	General Plan Module	103

7.4.3.1.1	General Plan Module Base Content	103

7.4.3.2	RT Prescription Module	103

7.4.3.2.1	RT Prescription Module Base Content	103

7.4.3.2.1.1	Referenced Standards	103

7.4.3.2.1.2	Module Definition	103

7.4.3.2.2	RT Prescription Module for Consistent Dose Tracking	103

7.4.3.3	RT Fraction Scheme Module	103

7.4.3.3.1	RT Fraction Scheme Module for Consistent Dose	103

7.4.3.3.2	RT Fraction Scheme Module for Delivery	103

7.4.3.3.2.1	Referenced Standards	103

7.4.3.3.2.2	Module Definition	103

7.4.4	Plan-Related Modules in Planning	104

7.4.4.1	Specific RT BeamType Specifications	104

7.4.4.1.1	RT Beams Module for Basic Static Beam	104

7.4.4.1.1.1	Referenced Standards	104

7.4.4.1.1.2	Module Definition	105

7.4.4.1.2	RT Beams Module for Basic Static MLC Beam	107

7.4.4.1.2.1	Referenced Standards	107

7.4.4.1.2.2	Module Definition	107

7.4.4.1.3	RT Beams Module for Arc Beam	109

7.4.4.1.3.1	Referenced Standards	109

7.4.4.1.3.2	Module Definition	109

7.4.4.1.4	RT Beams Module for MLC Fixed Aperture Arc Beam	111

7.4.4.1.4.1	Referenced Standards	111

7.4.4.1.4.2	Module Definition	111

7.4.4.1.5	RT Beams Module for MLC Variable Aperture Arc Beam	113

7.4.4.1.5.1	Referenced Standards	113

7.4.4.1.5.2	Module Definition	113

7.4.4.1.6	RT Beams Module for Hard Wedge Beam	114

7.4.4.1.6.1	Referenced Standards	114

7.4.4.1.6.2	Module Definition	115

7.4.4.1.7	RT Beams Module for Virtual Wedge Beam	117

7.4.4.1.7.1	Referenced Standards	117

7.4.4.1.7.2	Module Definition	117

7.4.4.1.8	RT Beams Module for Motorized Wedge Beam	119

7.4.4.1.8.1	Referenced Standards	119

7.4.4.1.8.2	Module Definition	120

7.4.4.1.9	RT Beams Module for Basic Static Electron Beam	122

7.4.4.1.9.1	Referenced Standards	122

7.4.4.1.9.2	Module Definition	122

7.4.4.1.10	RT Beams Module for Step & Shoot Beam	124

7.4.4.1.10.1	Referenced Standards	124

7.4.4.1.10.2	Module Definition	124

7.4.4.1.11	RT Beams Module for Sliding Window Beam	126

7.4.4.1.11.1	Referenced Standards	126

7.4.4.1.11.2	Module Definition	126

7.4.4.1.12	RT Beams Module for IMAT/VMAT Beam	128

7.4.4.1.12.1	Referenced Standards	128

7.4.4.1.12.2	Module Definition	129

7.4.4.1.13	RT Beams Module for Photon Applicator Beam	130

7.4.4.1.13.1	Referenced Standards	130

7.4.4.1.13.2	Module Definition	131

7.4.4.1.14	RT Beams Module for Photon Applicator Arc Beam	132

7.4.4.1.14.1	Referenced Standards	132

7.4.4.1.14.2	Module Definition	132

7.4.4.2	General Beam Attribute Specifications	135

7.4.4.2.1	Control Point Fixed Attribute List Base Content	135

7.4.4.2.1.1	Referenced Standards	135

7.4.4.2.1.2	Required Attributes	135

7.4.4.3	Beam Option Specifications	135

7.4.4.3.1	Bolus Beam Modifier Base Content	135

7.4.4.3.1.1	Referenced Standards	135

7.4.4.3.1.2	Required Attributes	136

7.4.4.3.2	Block Beam Modifier Base Content	136

7.4.4.3.2.1	Referenced Standards	136

7.4.4.3.2.2	Required Attributes	136

7.4.4.3.3	Compensator Beam Modifier Base Content	137

7.4.4.3.3.1	Referenced Standards	137

7.4.4.3.3.2	Required Attributes	137

7.4.4.3.4	Hard Wedge Beam Modifier Base Content	138

7.4.4.3.4.1	Referenced Standards	138

7.4.4.3.4.2	Required Attributes	138

# Introduction to this Supplement

This integration profile involves the exchange of RT Plan information between Treatment Planning Systems (TPS) and between Treatment Planning Systems and Treatment Management Systems (TMS). The emphasis for this profile is on reducing ambiguity involved in re-planning and incorporation of the planning information into the treatment management system in anticipation of transfer to a treatment delivery system. The transactions revolve around content rather than workflow.

This profile addresses a broad variety of “Beam Techniques” that exist in Radiation Therapy. Rather than define actors that had broad involvement in many optional transactions, a large number of actors were defined which have specific mandatory/required transactions and a small number of optional transactions related to beam modifiers. The actors are either producers or consumers of a DICOM RT Plan.

It is expected that the actual products commonly referred to as Treatment Planning Systems will implement one or more of the “producer” actors, and that the choice of which actors are implemented (for which adherence is claimed) will depend on the intended functionality (which is not defined by IHE-RO). A Treatment Planning System that is intended to be able to perform re-planning based on the output of another Treatment Planning System would be expected to adhere to one or more of the “consumer” actors.

It is expected that the actual products variously referred to as Oncology Information Systems, Oncology Information Management, or Electronic Medical Record for Oncology will implement the Treatment Management System Actor. While the profile does not dictate the functionality of the TMS, the TMS is responsible for providing an adequate view of the information provided to it (as a Beam Consumer) that in normal operating practice the appropriate user can ensure that the planning information has been properly consumed, associated with the correct patient, etc. As indicated in the table identifying actors and transactions, the TMS Actor must be able to act as the consumer in all retrieve transactions, i.e., consume all Beam Techniques.

## Open Issues and Questions

| # | Intr. in | Description |
| --- | --- | --- |

## Closed Issues

| # | Intr. in | Description |
| --- | --- | --- |
| 1 | Prepub 1.0 | This supplement is the successor of the ARTI supplement. The re-naming took place to be able to specify in more detail what this profile addresses as Treatment Planning – Plan Content was considered to be too broad and unspecific. |
| 2 | Prepub 1.0 | In addition to the re-naming of the profile the following changes were implemented (compared to the ARTI supplement): - The Primary Fluence Mode Sequence was added to all transactions. - High Dose Technique Type is now present in all transactions. - The Source to Wedge Tray Distance was removed from the Virtual and Motorized Wedge transactions. - All transactions for a Treatment Management System are optional. - The former (ARTI) actors for stereotactic treatments were re-named, as the term stereotactic is too overloaded and too unspecific. |
| 3 | Prepub 1.0 | Primary Fluence Mode Sequence: now required for all beams. If NON_STANDARD shall be of type “FFF”. This type is not defined by the Standard but it seems to be the de-facto-standard term. Should this be used or should we keep it open? What is about other terms, e.g., in case of “stereotactic” treatments? TC Decision (October 2013): Do not define specific IDs for non-standard cases. |
| 4 | Prepub 1.1 | Add CP 1138 changes (Average Beam Dose Parameters) to support the QAPV Profile. |
| 5 | Prepub 1.1 | High Dose Technique Type is of DICOM type 1C and type “NORMAL” has been removed from the DICOM standard by DICOM CP1132. Because of this CP the condition is now unambiguous. And therefore, for most of the beam types it now says that “If present, may not be ignored”. For all kinds of intensity-modulated beams as well as (former) stereotactic beams it now says “If present, shall be HDR and may not be ignored”. Is this what we wanted? TC 20140919: Already cleaned up in latest version of ARTI. |
| 6 | Prepub 1.0 | Any more notes on the new Average Beam Dose verification parameters in the Fraction Scheme? TC 20140919: done. |
| 7 | Prepub 1.3 | What’s about plans with mixed treatment modalities (e.g., beams and arcs)? Or is this a testing issue only? TC 20140919: Action Item for Bruce to draft a section in chapter 6.3. Done. |
| 8 | Prepub 1.4 | Re-consider the wording for the Table Top Positions as “shall be constant” could indicate that there is a value that is supposed to be constant. Perhaps “If present, shall be constant” (see proposal in B.3) TC 20141012: Attributes are O+* and shall state that if a value is present, it shall be constant. This does not force the type 2 value to become a Type 1. |
| 9 | Prepub 1.4 | Add attribute Beam Dose Meaning (300A,008B) as required with a value BEAM_LEVEL to the Fraction Group Sequence? TC 20141012: This is not the appropriate attribute to describe this. On the other hand it is well described in the DICOM Standard. |
| 10 | Prepub 1.4 | What is the actual use case for the Beam Dose at the Beam Dose specification point for a TPS consuming this value? (not questioning the use case of the Beam Doses for a TMS) TC 20141012: Adapted the requirements on the Type for producing and consuming actors to reflect actual clinical use cases |
| 11 | Prepub 1.4 | Block Tray ID? TC 20141012: Still R+, no change in requirement. |
| 12 | Prepub 1.4 | Number of Blocks for Electron Beams? Just 1? TC 20141012: For Electron Beams 0 or 1 blocks, photon beams unchanged. |
| 13 | Prepub 1.4 | Change requirement for SSD in Control Point Sequence: R+ (nor R+*) when Setup Technique is FIXED_SSD. TC 20141012: Updated the requirements with notation that was agreed upon for the Beam Dose issue and different actors in different use cases. |
| 14 | Prepub 1.4 | SSD is R+* for arcs? Is this correct? TC 20141013: Removed |
| 15 | Prepub 1.3 | Multi-target plans? TC 20140919: Action item for Uli and Jim to draft a proposal. TBD at the next face-to-face meeting. TC 20141013: Done, see the discussion for Beam Dose Representation Point. This is now also deferred to the Dose Consistency content profile. |
| 17 | Prepub 1.3 | Mixed treatment modalities with multiple treatment machines (Multi-Prescription plans)? TC 20140919: Action Item for Walter and Bruce to address the Clinical Advisory Sub-committee. TC 20141013: Decision that all beams within one plan are one the same machine. |
| 18 | Prepub 1.3 | Fluence Mode ID: is currently R+, which does not make much sense in case Fluence Mode is STANDARD. Proposal: remove attribute from profile as only DICOM condition should apply. TC 20141013: Introducing a new type “D” that does not change the DICOM type, but only requires the attribute value to be displayed. Action noted for Chris to CP the Technical Framework, including the new notion of <SCP>/<SCU> for attributes with different types for different actors. |
| 19 | Prepub 1.4 | Should we add CP 1331 “Treatment Time in RT Plan” to the profile as it provides an additional standardized channel for verification? (already added to the Fraction Group, to be deleted if not wanted). TC 20141013: Basic decision is to add it, but Uli and Jim to find out about overall delivery time (e.g., in case of gated treatments). 20141201: Answer from Jim: the CP basically reflects the internal behavior. 20150120: Rather a topic for TDPC, not for TPPC. |
| 20 | Prepub 1.4 | Feedback during testing: rename Conformal Arc and (MLC) Arc: in case the leaves are not changing call this a Static Conformal Arc. When the leaves are moving during gantry movement call this Dynamic Conformal (MLC) Arc. TC 20141012: rename Conformal Arc to MLC Variable Aperture Arc and MLC Arc to MLC Fixed Aperture Arc. |
| 21 | Prepub 1.4 | Discussion about usage of Beam Dose and Beam Dose Specification Point. TC 20141012: The TC agreed that it would be best to abandon the concept of the Beam Dose Specification Point but still there are systems in the field that rely on its presence. An updated approach was discussed where the Beam Dose is referenced through the Cumulative Dose Reference Coefficient back to the Dose Reference Item with Dose Reference Type TARGET. Chris to draft a proposal for a “dose consistency” profile for dose tracking where these parts will moved to and referenced from within this profile. TC 20141013: Based on the timeline for content profiles it needs to be decided if this still needs to be worked in to TPPC or if we stay with the separate profiles (background: when do we define the switch from ARTI to TPPC?) TC20150122: Decision to not address this for now, but rather develop the CDEB Profile and add new requirements regarding the Dose Reference Sequence in a next version of TPPC. |
| 22 | Prepub 1.4 | Add CP 1399 “Add Effective Wedge Angle” (in ballot). TC 20141012: Basically done: Added attribute to motorized wedge table, but DICOM tag is still missing – should be available soon for final text of the CP. 20141201: CP is in standard, attribute tags are updated. |
| 23 | Prepub 1.5 | 20141201: Add a TMS vs. TPS differentiation also for Referenced Dose Reference Sequence so that a TMS has to require these attributes, but a TPS does not necessarily have to consume this information,? 20150120: Decision to adapt it. See also item #26. 20150126: done |
| 24 | Prepub 1.5 | 20141201: Figure 6.5.2.2-1 direction of arrow of Beam Consumer to the archive: what does the arrow actually indicate? Data flow? “Basic interaction”? “Getting data from…”? 20150120: update it. 20150126: done |
| 26 | Prepub 1.7 | TC20150120: Update all consumer/producer condition pairs to producer/consumer. 20150126: done. |
| 25 | Prepub 1.6 | Update attribute of “Source to External Contour Distance” and “Average Beam Dose Point Source to External Contour Distance” once DICOM CP 1434 is ready. 20150120: waiting for March meeting of WG-06 20150504: Done |
| 27 | Prepub 1.7 | Results of discussion of TDPC: Add additional requirements for Series Date and Time to TPPC. Also add Equipment Module, Frame of Reference Module, 20150504: added referenced to TPPC Profile and re-using the definitions from there. |
| 28 | Prepub 1.9 | Issue 22 introduced the Effective Wedge Angle. The question is now, whether the second paragraph in chapter 3.31.4.1.2.1 is still required? TC20150510: Decision to keep paragraph for clarification |
| 29 | Prepub 1.9 | There are a few “left-over chapters” at the end of the document that have not yet been folded in the new DICOM Content Definition structure. How should these remains be best folded in as they do not represent a specific module, but parts of a module, referenced from multiple locations? In the beginning of the section on the Beams Module? Or the first chapter using these text macros, introducing them there as sub-chapters and then and referencing to there? TC20150510: Move to chapter 7.4.4.2 “General Beam Attribute Specifications” and 7.4.4.3 “Beam Option Specifications” |
| 16 | Prepub 1.4 | How to handle Plan Relationship of multiple DICOM RT Plans representing a conceptual treatment. Issue is kept open for Public Comment and WG-07 is asked to look at this question in parallel TC20180208: WG-07 recommends usage of term CONCURRENT in this case according to DICOM Standard. No additional requirement necessary for this profile |
| 17 | 1.2 (Trial) | Include CP 1658 + CP 1762 and CP 1659 |
| 18 | 1.2 (Trial) | TC20180122: Remove distances from profile (recommended by WG-07) – Beam Dose Verification Verification CP Sequence (300A,008C),Surface Entry Point (300A,012E), External Contour Entry Point (300A,0133), Source to Surface Distance (300A,0130), Source to External Contour Distance (300A,0132) |
| 19 | 1.2 (Trial) | Clarified wording Primary Fluence Mode – see issue #3 - Test instructions shall specify value for beam modelling |
| 20 | 1.2 (Trial) | TC20180208: Remove Beam Delivery Duration Limit (300A,00C5) |
| 21 | 1.2 (Trial) | TC20180208: Clarified specific rules for Source to Wedge Tray Distance (300A,00DA), Effective Wedge Angle(300A,00DE) and Wedge Angle (300A,00D5) for Virtual Wedge and Motorized Wedge |
| 22 | 1.3(Trial) | TC20180716: Remove Source to Surface Distance (300A,0130) and Source to External Contour Distance (300A,0132) from Control Point Sequence in accordance with WG-07 (Minutes 15th Jan, 2018) Adjusted order of producer/consumer description in Module Description of section 7.4 to match order in Types/Presence column |

Volume 1 – Profiles

## History of Annual Changes

Add the following to the IHE Technical Frameworks General Introduction Section 1.7:

Treatment Planning – Plan Content Profile:  This integration profile involves the exchange of RT Plan information between treatment planning systems and between treatment planning systems and treatment management systems. The emphasis for this profile is on reducing ambiguity involved in re-planning and incorporation of the planning information in to the treatment management system in anticipation of transfer to a treatment delivery system.

## Copyright Permission

Add the following to the IHE Technical Frameworks General Introduction Copyright section:

Add Section 6.

# Treatment Planning – Plan Content Integration (TPPC) Profile

This integration profile involves the exchange of RT Plan information between treatment planning systems and between treatment planning systems and treatment management systems. The emphasis for this profile is on reducing ambiguity involved in re-planning and incorporation of the planning information in to the treatment management system in anticipation of transfer to a treatment delivery system. The transactions revolve around content rather than workflow.

This profile addresses a broad variety of “Beam Techniques” that exist in Radiation Therapy. Rather than define actors that had broad involvement in many optional transactions, a large number of actors were defined which have specific mandatory/required transactions and a small number of optional transactions related to beam modifiers. The actors are either producers or consumers of a DICOM RT Plan.

## 6.1	TPPC Actors, Transactions, and Content Modules

Figure 6.1-1 shows the actors directly involved in the TPPC Profile and the relevant transactions between them. If needed for context, other actors that may be indirectly involved due to their participation in other related profiles are shown in dotted lines. Actors which have a mandatory grouping are shown in conjoined boxes.

TMS

Basic Static Beam
Producer

Archive

Basic Static Beam

Storage	Retrieval

Basic Static MLC Beam

Storage	Retrieval

Arc Beam

Storage	Retrieval

MLC Fixed Aperture Arc Beam

Storage	Retrieval

MLC Variable Aperture Arc Beam

Storage	Retrieval

Hard Wedge Beam

Storage	Retrieval

Virtual Wedge Beam

Storage	Retrieval

IMAT/VMAT Beam
Producer

Basic Static MLC Beam
Producer

Arc Beam
Consumer

Sliding Window Beam
Producer

Arc Beam
Producer

Motorized Wedge Beam
Producer

Photon Applicator Arc Beam Producer

Static Electron Beam
Producer

Hard Wedge Beam
Consumer

MLC Variable Aperture Arc Beam Consumer

Step & Shoot Beam
Producer

Photon Applicator Beam
Producer

Virtual Wedge Beam
Producer

Hard Wedge Beam
Producer

Motorized Wedge Beam
Consumer

Basic Static Beam
Consumer

Basic Static MLC Beam
Consumer

Step & Shoot Beam
Consumer

Static Electron Beam
Consumer

Virtual Wedge Beam
Consumer

MLC Fixed Aperture Arc Beam Consumer

IMAT/VMAT Beam
Consumer

Sliding Window Beam
Consumer

Photon Applicator Arc Beam Consumer

Photon Applicator Beam
Consumer

MLC Fixed Aperture Arc Beam Producer

MLC Variable Aperture Arc Beam Producer

Motorized Wedge Beam

Storage	Retrieval

Static Electron Beam

Storage	Retrieval

Step & Shoot Beam

Storage	Retrieval

Sliding Window Beam

Storage	Retrieval

IMAT/VMAT Beam

Storage	Retrieval

Stereotactic Beam

Storage	Retrieval

Stereotactic Arc Beam

Storage	Retrieval

Figure 6.1-1: TPPC Actor Diagram

Table 6.1-1 lists the transactions for each actor directly involved in the TPPC Profile. In order to claim support of this Profile, an implementation of an actor must perform the required transactions (labeled “R”) and may support the optional transactions (labeled “O”). Actor groupings are further described in Section 6.3.

Table 6.1-1: TPPC Profile - Actors and Transactions

| Actors | Transactions | Optionality | Section in Vol. 2 |
| --- | --- | --- | --- |
| Archive | Basic Static Beam Storage | R | TPPC-01 |
|  | Basic Static Beam Retrieval | R | TPPC-02 |
|  | Basic Static MLC Beam Storage | R | TPPC-03 |
|  | Basic Static MLC Beam Retrieval | R | TPPC-04 |
|  | Arc Beam Storage | R | TPPC-05 |
|  | Arc Beam Retrieval | R | TPPC-06 |
|  | MLC Fixed Aperture Arc Beam Storage | R | TPPC-07 |
|  | MLC Fixed Aperture Arc Beam Retrieval | R | TPPC-08 |
|  | MLC Variable Aperture Arc Beam Storage | R | TPPC-09 |
|  | MLC Variable Aperture Arc Beam Retrieval | R | TPPC-10 |
|  | Hard Wedge Beam Storage | R | TPPC-11 |
|  | Hard Wedge Beam Retrieval | R | TPPC-12 |
|  | Virtual Wedge Beam Storage | R | TPPC-13 |
|  | Virtual Wedge Beam Retrieval | R | TPPC-14 |
|  | Motorized Wedge Beam Storage | R | TPPC-15 |
|  | Motorized Wedge Beam Retrieval | R | TPPC-16 |
|  | Static Electron Beam Storage | R | TPPC-17 |
|  | Static Electron Beam Retrieval | R | TPPC-18 |
|  | Step & Shoot Beam Storage | R | TPPC-19 |
|  | Step & Shoot Beam Retrieval | R | TPPC-20 |
|  | Sliding Window Beam Storage | R | TPPC-21 |
|  | Sliding Window Beam Retrieval | R | TPPC-22 |
|  | IMAT/VMAT Beam Storage | R | TPPC-23 |
|  | IMAT/VMAT Beam Retrieval | R | TPPC-24 |
|  | Photon Applicator Beam Storage | R | TPPC-25 |
|  | Photon Applicator Beam Retrieval | R | TPPC-26 |
|  | Photon Applicator Arc Beam Storage | R | TPPC-27 |
|  | Photon Applicator Arc Beam Retrieval | R | TPPC-28 |
| Treatment Management System (TMS) (See Note Below) | Basic Static Beam Retrieval | O | TPPC-02 |
|  | Basic Static MLC Beam Retrieval | O | TPPC-04 |
|  | Arc Beam Retrieval | O | TPPC-06 |
|  | MLC Fixed Aperture Arc Beam Retrieval | O | TPPC-08 |
|  | MLC Variable Aperture Arc Beam Retrieval | O | TPPC-10 |
|  | Hard Wedge Beam Retrieval | O | TPPC-12 |
|  | Virtual Wedge Beam Retrieval | O | TPPC-14 |
|  | Motorized Wedge Beam Retrieval | O | TPPC-16 |
|  | Static Electron Beam Retrieval | O | TPPC-18 |
|  | Step & Shoot Beam Retrieval | O | TPPC-20 |
|  | Sliding Window Beam Retrieval | O | TPPC-22 |
|  | IMAT/VMAT Beam Retrieval | O | TPPC-24 |
|  | Photon Applicator Beam Retrieval | O | TPPC-26 |
|  | Photon Applicator Arc Beam Retrieval | O | TPPC-26 |
| Basic Static Beam Producer | Basic Static Beam Storage | R | TPPC-01 |
| Basic Static Beam Consumer | Basic Static Beam Retrieval | R | TPPC-02 |
| Basic Static MLC Beam Producer | Basic Static MLC Beam Storage | R | TPPC-03 |
| Basic Static MLC Beam Consumer | Basic Static MLC Beam Retrieval | R | TPPC-04 |
| Arc Beam Producer | Arc Beam Storage | R | TPPC-05 |
| Arc Beam Consumer | Arc Beam Retrieval | R | TPPC-06 |
| MLC Fixed Aperture Arc Beam Producer | MLC Fixed Aperture Arc Beam Storage | R | TPPC-07 |
| MLC Fixed Aperture Arc Beam Consumer | MLC Fixed Aperture Arc Beam Retrieval | R | TPPC-08 |
| MLC Variable Aperture Arc Beam Producer | MLC Variable Aperture Arc Beam Storage | R | TPPC-09 |
| MLC Variable Aperture Arc Beam Consumer | MLC Variable Aperture Arc Beam Retrieval | R | TPPC-10 |
| Hard Wedge Beam Producer | Hard Wedge Beam Storage | R | TPPC-11 |
| Hard Wedge Beam Consumer | Hard Wedge Beam Retrieval | R | TPPC-12 |
| Virtual Wedge Beam Producer | Virtual Wedge Beam Storage | R | TPPC-13 |
| Virtual Wedge Beam Consumer | Virtual Wedge Beam Retrieval | R | TPPC-14 |
| Motorized Wedge Beam Producer | Motorized Wedge Beam Storage | R | TPPC-15 |
| Motorized Wedge Beam Consumer | Motorized Wedge Beam Retrieval | R | TPPC-16 |
| Static Electron Beam Producer | Static Electron Beam Storage | R | TPPC-17 |
| Static Electron Beam Consumer | Static Electron Beam Retrieval | R | TPPC-18 |
| Step & Shoot Beam Producer | Step & Shoot Beam Storage | R | TPPC-19 |
| Step & Shoot Beam Consumer | Step & Shoot Beam Retrieval | R | TPPC-20 |
| Sliding Window Beam Producer | Sliding Window Beam Storage | R | TPPC-21 |
| Sliding Window Beam Consumer | Sliding Window Beam Retrieval | R | TPPC-22 |
| IMAT/VMAT Beam Producer | IMAT/VMAT Beam Storage | R | TPPC-23 |
| IMAT/VMAT Beam Consumer | IMAT/VMAT Beam Retrieval | R | TPPC-24 |
| Photon Applicator Beam Producer | Stereotactic Beam Storage | R | TPPC-25 |
| Photon Applicator Beam Consumer | Stereotactic Beam Retrieval | R | TPPC-26 |
| Photon Applicator Arc Beam Producer | Stereotactic Arc Beam Storage | R | TPPC-27 |
| Photon Applicator Arc Beam Consumer | Stereotactic Arc Beam Retrieval | R | TPPC-28 |

Note:  The TMS Integration Statement will indicate which transactions it is capable of supporting. In general, these will be grouped according to the overall functionality of the TMS Actor. For example, a general TMS would likely support all 14 transactions, while a Radiosurgery TMS may only support the Stereotactic-oriented beams. In addition, for cases where there are insufficient actors for complete testing of the TMS, the TMS can pass the Connectathon by claiming those transactions it successfully completed.

### 6.1.1 Actor Descriptions and Actor Profile Requirements

Normative requirements are typically documented in Volume 2 (Transactions) and Volume 3 (Content Modules). Some Integration Profiles, however, contain requirements which link transactions, data, and/or behavior. Those Profile requirements are documented in this section as normative requirements (“shall”).

## 6.2 TPPC Transaction Options

Options that may be selected for this Profile are listed in the Table 6.2-1 along with the Transactions to which they apply. In Table 6.1-1, each * Beam Producer has exactly one transaction, * Beam Storage. Similarly, each * Beam Consumer has exactly one transaction, * Beam Retrieval. For each of these, there are additional content options as noted in Table 5.2-1 from the set of {Bolus, Block, Compensator, Hard Wedge}. One or more of these content additions can be added to the base transaction based on the Integration Statement for the application. The Archive Actor must implement all options of all transactions. The TMS Actor must implement all options of all Retrieval transactions. Dependencies between options when applicable are specified in notes.

Table 6.2-1: TPPC - Transaction Options

| Transactions | Options | Optionality | Section in Vol 2 |
| --- | --- | --- | --- |
| Basic Static Beam Storage | Bolus Beam Modifier | O | 3.19 |
|  | Block Beam Modifier | O | 3.19 |
|  | Compensator Beam Modifier | O | 3.19 |
| Basic Static Beam Retrieval | Bolus Beam Modifier | O | 3.20 |
|  | Block Beam Modifier | O | 3.20 |
|  | Compensator Beam Modifier | O | 3.20 |
| Basic Static MLC Beam Storage | Bolus Beam Modifier | O | 3.21 |
|  | Compensator Beam Modifier | O | 3.21 |
| Basic Static MLC Beam Retrieval | Bolus Beam Modifier | O | 3.22 |
|  | Compensator Beam Modifier | O | 3.22 |
| Arc Beam Storage | Bolus Beam Modifier | O | 3.23 |
|  | Block Beam Modifier | O | 3.23 |
| Arc Beam Retrieval | Bolus Beam Modifier | O | 3.24 |
|  | Block Beam Modifier | O | 3.24 |
| MLC Fixed Aperture Arc Beam Storage | Bolus Beam Modifier | O | 3.25 |
| MLC Fixed Aperture Arc Beam Retrieval | Bolus Beam Modifier | O | 3.26 |
| MLC Variable Aperture Arc Beam Storage | Bolus Beam Modifier | O | 3.27 |
|  | Block Beam Modifier | O | 3.27 |
| MLC Variable Aperture Arc Beam Retrieval | Bolus Beam Modifier | O | 3.28 |
|  | Block Beam Modifier | O | 3.28 |
| Hard Wedge Beam Storage | Bolus Beam Modifier | O | 3.29 |
|  | Block Beam Modifier | O | 3.29 |
|  | Compensator Beam Modifier | O | 3.29 |
| Hard Wedge Beam Retrieval | Bolus Beam Modifier | O | 3.30 |
|  | Block Beam Modifier | O | 3.30 |
|  | Compensator Beam Modifier | O | 3.30 |
| Virtual Wedge Beam Storage | Bolus Beam Modifier | O | 3.31 |
|  | Block Beam Modifier | O | 3.31 |
|  | Compensator Beam Modifier | O | 3.31 |
|  | Hard Wedge Beam Modifier | O | 3.31 |
| Virtual Wedge Beam Retrieval | Bolus Beam Modifier | O | 3.32 |
|  | Block Beam Modifier | O | 3.32 |
|  | Compensator Beam Modifier | O | 3.32 |
|  | Hard Wedge Beam Modifier | O | 3.32 |
| Motorized Wedge Beam Storage | Bolus Beam Modifier | O | 3.33 |
|  | Block Beam Modifier | O | 3.33 |
|  | Compensator Beam Modifier | O | 3.33 |
|  | Hard Wedge Beam Modifier | O | 3.33 |
| Motorized Wedge Beam Retrieval | Bolus Beam Modifier | O | 3.34 |
|  | Block Beam Modifier | O | 3.34 |
|  | Compensator Beam Modifier | O | 3.34 |
|  | Hard Wedge Beam Modifier | O | 3.34 |
| Static Electron Beam Storage | Bolus Beam Modifier | O | 3.35 |
|  | Block Beam Modifier | O | 3.35 |
|  | Compensator Beam Modifier | O | 3.35 |
| Static Electron Beam Retrieval | Bolus Beam Modifier | O | 3.36 |
|  | Block Beam Modifier | O | 3.36 |
|  | Compensator Beam Modifier | O | 3.36 |
| Step & Shoot Beam Storage | Bolus Beam Modifier | O | 3.37 |
|  | Block Beam Modifier | O | 3.37 |
|  | Compensator Beam Modifier | O | 3.37 |
|  | Hard Wedge Beam Modifier | O | 3.37 |
| Step & Shoot Beam Retrieval | Bolus Beam Modifier | O | 3.38 |
|  | Block Beam Modifier | O | 3.38 |
|  | Compensator Beam Modifier | O | 3.38 |
|  | Hard Wedge Beam Modifier | O | 3.38 |
| Sliding Window Beam Storage | Bolus Beam Modifier | O | 3.39 |
|  | Block Beam Modifier | O | 3.39 |
|  | Compensator Beam Modifier | O | 3.39 |
|  | Hard Wedge Beam Modifier | O | 3.39 |
| Sliding Window Beam Retrieval | Bolus Beam Modifier | O | 3.40 |
|  | Block Beam Modifier | O | 3.40 |
|  | Compensator Beam Modifier | O | 3.40 |
|  | Hard Wedge Beam Modifier | O | 3.40 |
| IMAT/VMAT Beam Storage | Bolus Beam Modifier | O | 3.41 |
| IMAT/VMAT Beam Retrieval | Bolus Beam Modifier | O | 3.42 |
| Photon Applicator Beam Storage | Bolus Beam Modifier | O | 3.43 |
| Photon Applicator Beam Retrieval | Bolus Beam Modifier | O | 3.44 |
| Photon Applicator Arc Beam Storage | Bolus Beam Modifier | O | 3.45 |
| Photon Applicator Arc Beam Retrieval | Bolus Beam Modifier | O | 3.46 |

### 6.2.1	Producer / Consumer Transaction Groupings

Although not within the scope of this profile, it is possible for a clinical application to produce a treatment plan using several beams, each of which is defined by a different Beam Storage transaction as defined in this profile. For Producers, it is clinically acceptable for them to create a treatment plan with beams satisfying any of the Producer transactions with which they claim adherence. For such cases, a Consumer actor which also claims adherence to Beam Retrieval transactions encompassing the composited treatment plan may be able to receive such a plan containing a mixed transaction set of beams. However, if the Consumer actor cannot accept a plan with such a mixed set of beams, it must be able to handle the rejection of the plan in a safe manner.

Although an application may claim adherence to a set of transactions involving one or more Producer / Consumer pairs, the plan produced may not be deliverable by the equipment defined in a specific SOP instance created by a transaction. It is not required that all transactions defined in this profile are capable of being delivered by all delivery devices. As a result, there may be single transaction plans, or mixed transaction plans, that are not capable of being delivered by a given delivery device, even though it is defined as the delivery device in the treatment plan. The application user must be aware of the limitations of their delivery equipment, and should configure, where possible, their treatment planning systems and treatment management systems so that such incompatible plans cannot be created. Where such configuration is not possible, the application user should be aware of the limitations, and recognize that such plans may fail to be accepted when transferred to the delivery equipment, which is responsible for handling such incompatible plans in a safe manner.

## 6.3 TPPC Actor Required Groupings

Actor(s) which are required to be grouped with another actor(s) are listed in this section. The grouped actor may be from this profile or a different domain/profile. These mandatory required groupings, plus further descriptions if necessary, are given in the table below.

An actor from this profile (Column 1) must implement all of the required transactions in this profile in addition to all of the required transactions for the grouped profile/actor listed (Column 2).

Table 6.3-1: TPPC - Actors Required Groups

| TPPC Actor | Required Grouping Actor | Technical Framework Reference | Note |
| --- | --- | --- | --- |
| Archive | None | -- |  |
| Treatment Management System | None | -- |  |
| Basic Static Beam Producer | None | -- |  |
| Basic Static Beam Consumer | None | -- |  |
| Basic Static MLC Beam Producer | None | -- |  |
| Basic Static MLC Beam Consumer | None | -- |  |
| Arc Beam Producer | None | -- |  |
| Arc Beam Consumer | None | -- |  |
| MLC Fixed Aperture Arc Beam Producer | None | -- |  |
| MLC Fixed Aperture Arc Beam Consumer | None | -- |  |
| MLC Variable Aperture Arc Beam Producer | None | -- |  |
| MLC Variable Aperture Arc Beam Consumer | None | -- |  |
| Hard Wedge Beam Producer | None | -- |  |
| Hard Wedge Beam Consumer | None | -- |  |
| Virtual Wedge Beam Producer | None | -- |  |
| Virtual Wedge Beam Consumer | None | -- |  |
| Motorized Wedge Beam Producer | None | -- |  |
| Motorized Wedge Beam Consumer | None | -- |  |
| Static Electron Beam Producer | None | -- |  |
| Static Electron Beam Consumer | None | -- |  |
| Step & Shoot Beam Producer | None | -- |  |
| Step & Shoot Beam Consumer | None | -- |  |
| Sliding Window Beam Producer | None | -- |  |
| Sliding Window Beam Consumer | None | -- |  |
| IMAT/VMAT Beam Producer | None | -- |  |
| IMAT/VMAT Beam Consumer | None | -- |  |
| Photon Applicator Beam Producer | None | -- |  |
| Photon Applicator Beam Consumer | None | -- |  |
| Photon Applicator Arc Beam Producer | None | -- |  |
| Photon Applicator Arc Beam Consumer | None | -- |  |

## 6.4 TPPC Document Content Module

Not applicable.

## 6.5 TPPC Overview

This integration profile involves the exchange of RT Plan information between treatment planning systems and between treatment planning systems and treatment management systems. The emphasis for this profile is on reducing ambiguity involved in re-planning and incorporation of the planning information in to the treatment management system in anticipation of transfer to a treatment delivery system.

* Beam
Producer

Archive

* Beam
Consumer

* Beam Storage [TPPC-xx]

* Beam Retrieval [TPPC-yy]

There is one Producer actor and Storage transaction for each delivery technique in the profile.

There is one Consumer actor and Retrieval transaction for each delivery technique in the profile.

TMS

* Beam Retrieval [TPPC-yy]

The TMS has all transactions as optional. This allows the TMS Integration Statement to define its application scope.

* Transactions not numbered are placeholders for a transaction of that type per beam type.

Figure 6.5.1-1: Overall Process Flow in TPPC Profile

### 6.5.1 Concepts

This integration profile involves the exchange of RT Plan information between treatment planning systems and between treatment planning systems and treatment management systems. The emphasis for this profile is on reducing ambiguity involved in re-planning and incorporation of the planning information in to the treatment management system in anticipation of transfer to a treatment delivery system. The transactions revolve around content rather than workflow.

This profile addresses a broad variety of “Beam Techniques” that exist in Radiation Therapy. Rather than define actors that have broad involvement in many optional transactions, a large number of actors were defined which each have specific mandatory/required transactions and a small number of optional transactions related to beam modifiers. The actors are either producers or consumers of a DICOM RT Plan.

It is expected that the actual products commonly referred to as Treatment Planning Systems will implement one or more of the “producer” actors, and that the choice of which actors are implemented (for which adherence is claimed) will depend on the intended functionality (which is not defined by IHE-RO). A Treatment Planning System that is intended to be able to perform re-planning based on the output of another Treatment Planning System would be expected to adhere to one or more of the “consumer” actors.

It is expected that the actual products variously referred to as Oncology Information Systems, Oncology Information Management, or Electronic Medical Record for Oncology will implement the Treatment Management System (TMS) Actor. While the profile does not dictate the functionality of the TMS, the TMS is responsible for providing an adequate view of the information provided to it (as a Beam Consumer) such that, in normal operating practice, the appropriate user can ensure that the planning information has been properly consumed, associated with the correct patient, etc. No transactions have been defined between the TMS Actor in this profile and the TMS Actor in other profiles, and any necessary interface is considered private (in the same way that an Image Manager and an Image Archive are related in the Radiology Domain Scheduled Workflow Profile). In practice, it is expected that once a TMS has consumed the information provided to it by a Beam Producer, the system incorporating the TMS Actor will then be able to act as the TMS in delivery-oriented profiles and provide that information to a Treatment Delivery System Actor in that profile. It is not expected that a TMS Actor for this profile from one vendor will interoperate with a TMS Actor for other delivery profiles from another vendor. As indicated in the table identifying actors and transactions, the TMS Actor can support retrieval of any of the beam types (all transactions are optional). The TMS shall indicate in its Integration Statement the scope of its capabilities (i.e., which beam types it supports). It is expected that a TMS will support most, if not all, beam types. However, there may be beam types for which full testing is not possible due to limitations on the number of producers of a specific beam type, hence the optional transaction list.

It should also be noted that chapter 7 in this Supplement’s Volume 3 specifies content that is mandatory across all transactions.

Finally, there are individual attributes within a RT Plan that are not specified in this profile, but have significant safety implications if ignored. As much as possible, these attributes have been identified in the transactions and it is indicated that a ‘retrieval’ actor shall handle RT Plans that may include these attributes in a safe manner. This behavior can include rejection of the RT Plan or appropriate warmings (with user acknowledgement) as possible courses of action in such circumstances.

### 6.5.2 Use Case #1: Treatment Replanning

A radiation oncology treatment planning system (TPS) creates and stores a treatment plan according to one of the techniques specified. A second TPS, requiring the initial plan in order to add/change information, retrieves the original treatment plan, modifies the plan, and then stores the modified treatment plan as a new plan to the archive.

#### 6.5.2.1 Treatment Re-planning Use Case Description

A TPS, through transactions not described in this profile (but similar to the IHE-RO Basic Radiation Therapy Objects Profile) acquires sufficient information to create a treatment plan according to one of the treatment techniques described in this profile.

The treatment plan is stored to the archive.

A second TPS, required the original treatment plan in order to modify it, will read in the original plan, perform modifications to the plan, and then save the modified plan as a new plan to the archive.

#### 6.5.2.2 Treatment Replanning Process Flow

* Beam
Producer

Archive

* Beam
Consumer

* Beam Storage [TPPC-xx]

There is one Producer actor and Storage transaction for each delivery technique in the profile.

There is one Consumer actor and Retrieval transaction for each delivery technique in the profile.

* Beam Retrieval [TPPC-yy]

* Transactions not numbered are placeholders for a transaction of that type per beam type.

Figure 6.5.2.2-1: Process Flow for Treatment Replanning Use Case

### 6.5.3 Use Case #2: TMS Upload

A radiation oncology treatment planning system (TPS) creates and stores a treatment plan according to one of the techniques specified. The Treatment Management System (TMS) retrieves the treatment plan in order to process it for further use and potential delivery (not included in this profile).

#### 6.5.3.1 TMS Upload Use Case Description

A TPS, through transactions not described in this profile (but similar to the IHE-RO Basic Radiation Therapy Objects Profile) acquires sufficient information to create a treatment plan according to one of the treatment techniques described in this profile.

The treatment plan is stored to the archive.

A TMS retrieves the plan from the archive.

#### 6.5.3.2 TMS Upload Process Flow

* Beam
Producer

Archive

* Beam Storage [TPPC-xx]

There is one Producer actor and Storage transaction for each delivery technique in the profile.

TMS

* Beam Retrieval [TPPC-yy]

The TMS has all transactions as optional. This allows the TMS Integration Statement to define its application scope.

* Transactions not numbered are placeholders for a transaction of that type per beam type.

Figure 6.5.3.2-1: Process Flow for TMS Upload Use Case

## TPPC Security Considerations

Not Applicable

## TPPC Cross Profile Considerations

Not Applicable

Appendices

Appendix A Actor Summary Definitions

Add the following terms to the IHE TF General Introduction Appendix A:

Basic Static Beam Producer – A Treatment Planning System (TPS) capable of producing a radiation therapy treatment plan with static, non-MLC, treatment beams.

Basic Static Beam Consumer – A Treatment Planning System (TPS) capable of consuming a radiation therapy treatment plan with static, non-MLC, treatment beams.

Basic Static MLC Beam Producer – A Treatment Planning System (TPS) capable of producing a radiation therapy treatment plan with static, MLC, treatment beams.

Basic Static MLC Beam Consumer – A Treatment Planning System (TPS) capable of consuming a radiation therapy treatment plan with static, MLC, treatment beams.

Arc Beam Producer – A Treatment Planning System (TPS) capable of producing a radiation therapy treatment plan with non-MLC Fixed Aperture Arc treatment beams.

Arc Beam Consumer – A Treatment Planning System (TPS) capable of consuming a radiation therapy treatment plan with non-MLC Fixed Aperture Arc treatment beams.

MLC Fixed Aperture Arc Beam Producer – A Treatment Planning System (TPS) capable of producing a radiation therapy treatment plan with MLC Fixed Aperture Arc treatment beams.

MLC Fixed Aperture Arc Beam Consumer – A Treatment Planning System (TPS) capable of consuming a radiation therapy treatment plan with MLC Fixed Aperture Arc treatment beams.

MLC Variable Aperture Arc Beam Producer – A Treatment Planning System (TPS) capable of producing a radiation therapy treatment plan with MLC Variable Aperture Arc treatment beams.

MLC Variable Aperture Arc Beam Consumer – A Treatment Planning System (TPS) capable of consuming a radiation therapy treatment plan with MLC Variable Aperture Arc treatment beams.

Hard Wedge Beam Producer – A Treatment Planning System (TPS) capable of producing a radiation therapy treatment plan with static treatment beams using physical wedges.

Hard Wedge Beam Consumer – A Treatment Planning System (TPS) capable of consuming a radiation therapy treatment plan with static treatment beams using physical wedges.

Virtual Wedge Beam Producer – A Treatment Planning System (TPS) capable of producing a radiation therapy treatment plan with static treatment beams using virtual wedges.

Virtual Wedge Beam Consumer – A Treatment Planning System (TPS) capable of consuming a radiation therapy treatment plan with static treatment beams using virtual wedges.

Motorized Wedge Beam Producer – A Treatment Planning System (TPS) capable of producing a radiation therapy treatment plan with static treatment beams using motorized wedges.

Motorized Wedge Beam Consumer – A Treatment Planning System (TPS) capable of consuming a radiation therapy treatment plan with static treatment beams using motorized wedges.

Static Electron Beam Producer – A Treatment Planning System (TPS) capable of producing a radiation therapy treatment plan with static electron treatment beams.

Static Electron Beam Consumer – A Treatment Planning System (TPS) capable of consuming a radiation therapy treatment plan with static electron treatment beams.

Step & Shoot Beam Producer – A Treatment Planning System (TPS) capable of producing a radiation therapy treatment plan with step & shoot IMRT treatment beams.

Step & Shoot Beam Consumer – A Treatment Planning System (TPS) capable of consuming a radiation therapy treatment plan with step & shoot IMRT treatment beams.

Sliding Window Beam Producer – A Treatment Planning System (TPS) capable of producing a radiation therapy treatment plan with sliding window IMRT treatment beams.

Sliding Window Beam Consumer – A Treatment Planning System (TPS) capable of consuming a radiation therapy treatment plan with sliding window IMRT treatment beams.

IMAT/VMAT Beam Producer – A Treatment Planning System (TPS) capable of producing a radiation therapy treatment plan with VMAT/IMAT IMRT treatment beams.

IMAT/VMAT Beam Consumer – A Treatment Planning System (TPS) capable of consuming a radiation therapy treatment plan with VMAT/IMAT IMRT treatment beams.

Photon Applicator Beam Producer – A Treatment Planning System (TPS) capable of producing a radiation therapy treatment plan with static, stereotactic treatment beams.

Photon Applicator Beam Consumer – A Treatment Planning System (TPS) capable of consuming a radiation therapy treatment plan with static, stereotactic treatment beams.

Photon Applicator Arc Beam Producer – A Treatment Planning System (TPS) capable of producing a radiation therapy treatment plan with stereotactic arc treatment beams.

Photon Applicator Arc Beam Consumer – A Treatment Planning System (TPS) capable of consuming a radiation therapy treatment plan with stereotactic arc treatment beams.

Treatment Management System (TMS) – An application providing radiation oncology management services and capable of consuming treatment plans with any of the above treatment techniques.

Transaction Summary Definitions

Add the following terms to the IHE TF General Introduction Namespace list of Transactions:

TPPC-01:  Basic Static Beam Storage

In the Basic Static Beam Storage transaction, a Static Beam Producer stores a treatment plan to the Archive. The treatment plan shall contain only static, non-MLC treatment beams.

TPPC-02:  Basic Static Beam Retrieval

In the Basic Static Beam Retrieval transaction, a Static Beam Consumer or a TMS receives a treatment plan from the Archive. The treatment plan shall contain only static, non-MLC treatment beams.

TPPC-03:  Basic Static MLC Beam Storage

In the Basic Static MLC Beam Storage transaction, a Static MLC Beam Producer stores a treatment plan to the Archive. The treatment plan shall contain only static, MLC treatment beams.

TPPC-04:  Basic Static MLC Beam Retrieval

In the Basic Static MLC Beam Retrieval transaction, a Static MLC Beam Consumer or a TMS receives a treatment plan from the Archive. The treatment plan shall contain only static, MLC treatment beams.

TPPC-05:  Arc Beam Storage

In the Arc Beam Storage transaction, an Arc Beam Producer stores a treatment plan to the Archive. The treatment plan shall contain only non-MLC Fixed Aperture Arc treatment beams.

TPPC-06:  Arc Beam Retrieval

In the Arc Beam Retrieval transaction, an Arc Beam Consumer or a TMS receives a treatment plan from the Archive. The treatment plan shall contain only non-MLC Fixed Aperture Arc treatment beams.

TPPC-07:  MLC Fixed Aperture Arc Beam Storage

In the MLC Fixed Aperture Arc Beam Storage transaction, an MLC Fixed Aperture Arc Beam Producer stores a treatment plan to the Archive. The treatment plan shall contain only MLC Fixed Aperture Arc treatment beams.

TPPC-08:  MLC Fixed Aperture Arc Beam Retrieval

In the MLC Fixed Aperture Arc Beam Retrieval transaction, an MLC Fixed Aperture Arc Beam Consumer or a TMS receives a treatment plan from the Archive. The treatment plan shall contain only MLC Fixed Aperture Arc treatment beams.

TPPC-09:  MLC Variable Aperture Arc Beam Storage

In the MLC Variable Aperture Arc Beam Storage transaction, a MLC Variable Aperture Arc Beam Producer stores a treatment plan to the Archive. The treatment plan shall contain only MLC Variable Aperture Arc treatment beams.

TPPC-10:  MLC Variable Aperture Arc Beam Retrieval

In the MLC Variable Aperture Arc Beam Retrieval transaction, an MLC Variable Aperture Arc Beam Consumer or a TMS receives a treatment plan from the Archive. The treatment plan shall contain only MLC Variable Aperture Arc treatment beams.

TPPC-11:  Hard Wedge Beam Storage

In the Hard Wedge Beam Storage transaction, a Hard Wedge Beam Producer stores a treatment plan to the Archive. The treatment plan shall contain only static treatment beams using physical wedges.

TPPC-12:  Hard Wedge Beam Retrieval

In the Hard Wedge Beam Retrieval transaction, a Hard Wedge Beam Consumer or a TMS receives a treatment plan from the Archive. The treatment plan shall contain only static treatment beams using physical wedges.

TPPC-13:  Virtual Wedge Beam Storage

In the Virtual Wedge Beam Storage transaction, a Virtual Wedge Beam Producer stores a treatment plan to the Archive. The treatment plan shall contain only static treatment beams using virtual wedges.

TPPC-14:  Virtual Wedge Beam Retrieval

In the Virtual Wedge Beam Retrieval transaction, a Virtual Wedge Beam Consumer or a TMS receives a treatment plan from the Archive. The treatment plan shall contain only static treatment beams using virtual wedges.

TPPC-15:  Motorized Wedge Beam Storage

In the Motorized Wedge Beam Storage transaction, a Motorized Wedge Beam Producer stores a treatment plan to the Archive. The treatment plan shall contain only static treatment beams using motorized wedges.

TPPC-16:  Motorized Wedge Beam Retrieval

In the Motorized Wedge Beam Retrieval transaction, a Motorized Wedge Beam Consumer or a TMS receives a treatment plan from the Archive. The treatment plan shall contain only static treatment beams using motorized wedges.

TPPC-17:  Static Electron Beam Storage

In the Static Electron Beam Storage transaction, a Static Electron Beam Producer stores a treatment plan to the Archive. The treatment plan shall contain only static electron treatment beams.

TPPC-18:  Static Electron Beam Retrieval

In the Static Electron Beam Retrieval transaction, a Static Electron Beam Consumer or a TMS receives a treatment plan from the Archive. The treatment plan shall contain only static electron treatment beams.

TPPC-19:  Step & Shoot Beam Storage

In the Step & Shoot Beam Storage transaction, a Step & Shoot Beam Producer stores a treatment plan to the Archive. The treatment plan shall contain only step & shoot IMRT treatment beams.

TPPC-20:  Step & Shoot Beam Retrieval

In the Step & Shoot Beam Retrieval transaction, a Step & Shoot Beam Consumer or a TMS receives a treatment plan from the Archive. The treatment plan shall contain only step & shoot IMRT treatment beams.

TPPC-21:  Sliding Window Beam Storage

In the Sliding Window Beam Storage transaction, a Sliding Window Beam Producer stores a treatment plan to the Archive. The treatment plan shall contain only sliding window IMRT treatment beams.

TPPC-22:  Sliding Window Beam Retrieval

In the Sliding Window Beam Retrieval transaction, a Sliding Window Beam Consumer or a TMS receives a treatment plan from the Archive. The treatment plan shall contain only sliding window IMRT treatment beams.

TPPC-23:  IMAT/VMAT Beam Storage

In the IMAT/VMAT Beam Storage transaction, an IMAT/VMAT Beam Producer stores a treatment plan to the Archive. The treatment plan shall contain only IMAT/VMAT IMRT treatment beams.

TPPC-24:  IMAT/VMAT Beam Retrieval

In the IMAT/VMAT Beam Retrieval transaction, an IMAT/VMAT Beam Consumer or a TMS receives a treatment plan from the Archive. The treatment plan shall contain only IMAT/VMAT IMRT treatment beams.

TPPC-25:  Photon Applicator Beam Storage

In the Photon Applicator Beam Storage transaction, a Photon Applicator Beam Producer stores a treatment plan to the Archive. The treatment plan shall contain only static treatment beams using photon applicators.

TPPC-26:  Photon Applicator Beam Retrieval

In the Photon Applicator Beam Retrieval transaction, a Photon Applicator Beam Consumer or a TMS receives a treatment plan from the Archive. The treatment plan shall contain only static treatment beams using photon applicator.

TPPC-27:  Photon Applicator Arc Beam Storage

In the Stereotactic Arc Beam Storage transaction, a Stereotactic Arc Beam Producer stores a treatment plan to the Archive. The treatment plan shall contain only stereotactic arc treatment beams.

TPPC-28:  Photon Applicator Arc Beam Retrieval

In the Stereotactic Arc Beam Retrieval transaction, a Stereotactic Arc Beam Consumer or a TMS receives a treatment plan from the Archive. The treatment plan shall contain only stereotactic arc treatment beams.

Glossary

Add the following terms to the IHE Technical Frameworks General Introduction Glossary:

None

Volume 2 – Transactions

Add Section 3.19.

## 3.19 TPPC-01 Basic Static Beam Storage

### 3.19.1 Scope

In the Basic Static Beam Storage transaction, a Producer of an RT Plan that incorporates the beam technique identified in TPPC-01: Basic Static Beam Storage stores the plan to the archive

### 3.19.2 Use Case Roles

Basic Static Beam Storage

Archive

Basic Static Beam Producer

| Actor: | Basic Static Beam Producer |
| --- | --- |
| Role: | Creates Basic Static Beam RT Plan and stores plan to an RT Archive |
| Actor: | Archive |
| Role: | Accept and store RT Plan from Basic Static Beam Producer |

### 3.19.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.19.4 Interaction Diagram

Archive

Basic Static Beam Producer

C-STORE (RT Plan)

#### 3.19.4.1 Basic Static Beam Storage

##### 3.19.4.1.1 Trigger Events

The Basic Static Beam Producer transfers the plan to the Archive once the plan is created and the dose calculation is finished.

##### 3.19.4.1.2 Message Semantics

The Basic Static Beam Producer uses the DICOM C-STORE message to transfer the plan.

The Basic Static Beam Producer is the DICOM Storage SCU and the Archive is the DICOM Storage SCP.

The Basic Static Beam Producer may create a new series containing the plan or may use an existing series, where previous plan(s) are contained.

The study where the series of the plan is contained shall be the same study as the one containing the structure set referenced in the plan.

###### 3.19.4.1.2.1 Storage of RT Plan containing a Basic Static Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.19.4.1.2.2 Optional Modifiers

The Basic Static Beam Producer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Compensator Beam Modifier | 7.4.4.3.3 |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.19.4.1.3 Expected Actions

The Archive stores the RT Plan.

### 3.19.5 Security Considerations

There are no specific security considerations.

## 3.20 TPPC-02: Basic Static Beam Retrieval

### 3.20.1 Scope

In the Basic Static Beam Retrieval transaction, a consumer of an RT Plan that incorporates the beam technique identified in TPPC-01: Basic Static Beam Storage, retrieves the plan from the archive.

### 3.20.2 Use Case Roles

Basic Static Beam Retrieval

Archive

Basic Static Beam Consumer

| Actor: | Basic Static Beam Consumer |
| --- | --- |
| Role: | Stores plan transmitted from Archive |
| Actor: | Archive |
| Role: | Transmits Plan to Basic Static Beam Consumer |

### 3.20.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.20.4 Interaction Diagram

Archive

Basic Static Beam Consumer

C_STORE (RT Plan)

#### 3.20.4.1 Basic Static Beam Retrieval

##### 3.20.4.1.1 Trigger Events

The Archive transfers the plan to the Basic Static Beam Consumer.

##### 3.20.4.1.2 Message Semantics

The Archive uses the DICOM C-STORE message to transfer the plan.

The Archive is the DICOM Storage SCU and the Basic Static Beam Consumer is the DICOM Storage SCP.

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.20.4.1.2.1 Storage of RT Plan containing a Basic Static Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.20.4.1.2.2 Optional Modifiers

The Basic Static Beam Consumer may support the following optional modifications :

| Optional Modifiers | Section |
| --- | --- |
| Compensator Beam Modifier | 7.4.4.3.3 |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.20.4.1.3 Expected Actions

The Basic Static Beam Consumer stores the RT Plan.

### 3.20.5 Security Considerations

There are no specific security considerations.

## 3.21 TPPC-03 Basic Static MLC Beam Storage

### 3.21.1 Scope

In the Basic Static MLC Beam Storage transaction, a Producer of an RT Plan that incorporates the beam technique identified in TPPC-03: Basic Static MLC Beam Storage stores the plan to the archive

### 3.21.2 Use Case Roles

Basic Static MLC Beam Storage

Archive

Basic Static MLC Beam Producer

| Actor: | Basic Static MLC Beam Producer |
| --- | --- |
| Role: | Creates Basic Static MLC Beam RT Plan and stores plan to an RT Archive |
| Actor: | Archive |
| Role: | Accept and store RT Plan from Basic Static MLCBeam Producer |

### 3.21.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.21.4 Interaction Diagram

Archive

Basic Static MLC Beam Producer

C-STORE (RT Plan)

#### 3.21.4.1 Basic Static MLC Beam Storage

##### 3.21.4.1.1 Trigger Events

The Basic Static MLC Beam Producer transfers the plan to the Archive once the plan is created and the dose calculation is finished.

##### 3.21.4.1.2 Message Semantics

The Basic Static MLC Beam Producer uses the DICOM C-STORE message to transfer the plan.

The Basic Static MLC Beam Producer is the DICOM Storage SCU and the Archive is the DICOM Storage SCP.

The Basic Static MLC Beam Producer may create a new series containing the plan or may use an existing series, where previous plan(s) are contained.

The study where the series of the plan is contained shall be the same study as the one containing the structure set referenced in the plan.

###### 3.21.4.1.2.1 Storage of RT Plan containing a Basic Static MLC Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.21.4.1.2.2 Optional Modifiers

The Basic Static MLC Beam Producer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Compensator Beam Modifier | 7.4.4.3.3 |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.21.4.1.3 Expected Actions

The Archive stores the RT Plan.

### 3.21.5 Security Considerations

There are no specific security considerations.

## 3.22 TPPC-04: Basic Static MLC Beam Retrieval

### 3.22.1 Scope

In the Basic Static MLC Beam Retrieval transaction, a consumer of an RT Plan that incorporates the beam technique identified in TPPC-03: Basic Static MLC Beam Storage, retrieves the plan from the archive.

### 3.22.2 Use Case Roles

Basic Static MLC Beam Retrieval

Archive

Basic Static MLC Beam Consumer

| Actor: | Basic Static MLC Beam Consumer |
| --- | --- |
| Role: | Stores plan transmitted from Archive |
| Actor: | Archive |
| Role: | Transmits Plan to Basic Static MLC Beam Consumer |

### 3.22.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.22.4 Interaction Diagram

Archive

Basic Static MLC Beam Consumer

C_STORE (RT Plan)

#### 3.22.4.1 Basic Static MLC Beam Retrieval

##### 3.22.4.1.1 Trigger Events

The Archive transfers the plan to the Basic Static Beam Consumer.

##### 3.22.4.1.2 Message Semantics

The Archive uses the DICOM C-STORE message to transfer the plan.

The Archive is the DICOM Storage SCU and the Basic Static MLC Beam Consumer is the DICOM Storage SCP.

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.22.4.1.2.1 Storage of RT Plan containing a Basic Static MLC Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.22.4.1.2.2 Optional Modifiers

The Basic Static Beam Consumer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Compensator Beam Modifier | 7.4.4.3.3 |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.22.4.1.3 Expected Actions

The Basic Static MLC Beam Consumer stores the RT Plan.

### 3.22.5 Security Considerations

There are no specific security considerations.

## 3.23 TPPC-05 Arc Beam Storage

### 3.23.1 Scope

In the Arc Beam Storage transaction, a Producer of an RT Plan that incorporates the beam technique identified in TPPC-05: Arc Beam Storage stores the plan to the archive

### 3.23.2 Use Case Roles

Arc Beam Storage

Archive

Arc Beam Producer

| Actor: | Arc Beam Producer |
| --- | --- |
| Role: | Creates Arc Beam RT Plan and stores plan to an RT Archive |
| Actor: | Archive |
| Role: | Accept and store RT Plan from Basic Arc Beam Producer |

### 3.23.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.23.4 Interaction Diagram

Archive

Arc Beam Producer

C-STORE (RT Plan)

#### 3.23.4.1 Arc Beam Storage

##### 3.23.4.1.1 Trigger Events

The Arc Beam Producer transfers the plan to the Archive once the plan is created and the dose calculation is finished.

##### 3.23.4.1.2 Message Semantics

The Arc Beam Producer uses the DICOM C-STORE message to transfer the plan.

The Arc Beam Producer is the DICOM Storage SCU and the Archive is the DICOM Storage SCP.

The Arc Beam Producer may create a new series containing the plan or may use an existing series, where previous plan(s) are contained.

The study where the series of the plan is contained shall be the same study as the one containing the structure set referenced in the plan.

###### 3.23.4.1.2.1 Storage of RT Plan containing an Arc Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.23.4.1.2.2 Optional Modifiers

The Arc Beam Producer may support the following optional :

| Optional Modifiers | Section |
| --- | --- |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.23.4.1.3 Expected Actions

The Archive stores the RT Plan.

### 3.23.5 Security Considerations

There are no specific security considerations.

## 3.24 TPPC-06: Arc Beam Retrieval

### 3.24.1 Scope

In the Arc Beam Retrieval transaction, a consumer of an RT Plan that incorporates the beam technique identified in TPPC-05: Arc Beam Storage, retrieves the plan from the archive.

### 3.24.2 Use Case Roles

Arc Beam Retrieval

Archive

Arc Beam Consumer

| Actor: | Arc Beam Consumer |
| --- | --- |
| Role: | Stores plan transmitted from Archive |
| Actor: | Archive |
| Role: | Transmits Plan to Arc Beam Consumer |

### 3.24.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.24.4 Interaction Diagram

Archive

Arc Beam Consumer

C_STORE (RT Plan)

#### 3.24.4.1 Arc Beam Retrieval

##### 3.24.4.1.1 Trigger Events

The Archive transfers the plan to the Arc Beam Consumer.

##### 3.24.4.1.2 Message Semantics

The Archive uses the DICOM C-STORE message to transfer the plan.

The Archive is the DICOM Storage SCU and the Arc Beam Consumer is the DICOM Storage SCP.

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.24.4.1.2.1 Storage of RT Plan containing an Arc Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.24.4.1.2.2 Optional Modifiers

The Arc Beam Consumer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.24.4.1.3 Expected Actions

The Arc Beam Consumer stores the RT Plan.

### 3.24.5 Security Considerations

There are no specific security considerations.

## 3.25 TPPC-07 MLC Fixed Aperture Arc Beam Storage

### 3.25.1 Scope

In the MLC Fixed Aperture Arc Beam Storage transaction, a Producer of an RT Plan that incorporates the beam technique identified in TPPC-07: MLC Fixed Aperture Arc Beam Storage stores the plan to the archive

### 3.25.2 Use Case Roles

MLC Fixed Aperture Arc Beam Storage

Archive

MLC Fixed Aperture Arc Beam Producer

| Actor: | MLC Fixed Aperture Arc Beam Producer |
| --- | --- |
| Role: | Creates MLC Fixed Aperture Arc Beam RT Plan and stores plan to an RT Archive |
| Actor: | Archive |
| Role: | Accept and store RT Plan from MLC Fixed Aperture Arc Beam Producer |

### 3.25.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.25.4 Interaction Diagram

Archive

MLC Fixed Aperture Arc Beam Producer

C-STORE (RT Plan)

#### 3.25.4.1 MLC Fixed Aperture Arc Beam Storage

##### 3.25.4.1.1 Trigger Events

The MLC Fixed Aperture Arc Beam Producer transfers the plan to the Archive once the plan is created and the dose calculation is finished.

##### 3.25.4.1.2 Message Semantics

The MLC Fixed Aperture Arc Beam Producer uses the DICOM C-STORE message to transfer the plan.

The MLC Fixed Aperture Arc Beam Producer is the DICOM Storage SCU and the Archive is the DICOM Storage SCP.

The MLC Fixed Aperture Arc Beam Producer may create a new series containing the plan or may use an existing series, where previous plan(s) are contained.

The study where the series of the plan is contained shall be the same study as the one containing the structure set referenced in the plan.

###### 3.25.4.1.2.1 Storage of RT Plan containing a MLC Fixed Aperture Arc Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.25.4.1.2.2 Optional Modifiers

The MLC Fixed Aperture Arc Beam Producer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Bolus Beam Modifier | 7.4.4.3.1 |

##### 3.25.4.1.3 Expected Actions

The Archive stores the RT Plan.

### 3.25.5 Security Considerations

There are no specific security considerations.

## 3.26 TPPC-08: MLC FIXED APERTURE ARC Beam Retrieval

### 3.26.1 Scope

In the MLC Fixed Aperture Arc Beam Retrieval transaction, a consumer of an RT Plan that incorporates the beam technique identified in TPPC-07: MLC Fixed Aperture Arc Beam Storage, retrieves the plan from the archive.

### 3.26.2 Use Case Roles

MLC Fixed Aperture Arc Beam Retrieval

Archive

MLC Fixed Aperture Arc Beam Consumer

| Actor: | MLC Fixed Aperture Arc Beam Consumer |
| --- | --- |
| Role: | Stores plan transmitted from Archive |
| Actor: | Archive |
| Role: | Transmits Plan to MLC Fixed Aperture Arc Beam Consumer |

### 3.26.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.26.4 Interaction Diagram

Archive

MLC Fixed Aperture Arc Beam Consumer

C_STORE (RT Plan)

#### 3.26.4.1 MLC Fixed Aperture Arc Beam Retrieval

##### 3.26.4.1.1 Trigger Events

The Archive transfers the plan to the MLC Fixed Aperture Arc Beam Consumer.

##### 3.26.4.1.2 Message Semantics

The Archive uses the DICOM C-STORE message to transfer the plan.

The Archive is the DICOM Storage SCU and the MLC Fixed Aperture Arc Beam Consumer is the DICOM Storage SCP.

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.26.4.1.2.1 Storage of RT Plan containing a MLC Fixed Aperture Arc Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.26.4.1.2.2 Optional Modifiers

The MLC Fixed Aperture Arc Beam Consumer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Bolus Beam Modifier | 7.4.4.3.1 |

##### 3.26.4.1.3 Expected Actions

The MLC Fixed Aperture Arc Beam Consumer stores the RT Plan.

### 3.26.5 Security Considerations

There are no specific security considerations.

## 3.27 TPPC-09 MLC Variable Aperture Arc Beam Storage

### 3.27.1 Scope

In the MLC Variable Aperture Arc Beam Storage transaction, a Producer of an RT Plan that incorporates the beam technique identified in TPPC-09: MLC Variable Aperture Arc Beam Storage stores the plan to the archive

### 3.27.2 Use Case Roles

MLC Variable Aperture Arc Beam Storage

Archive

MLC Variable Aperture Arc Beam Producer

| Actor: | MLC Variable Aperture Arc Beam Producer |
| --- | --- |
| Role: | Creates MLC Variable Aperture Arc Beam RT Plan and stores plan to an RT Archive |
| Actor: | Archive |
| Role: | Accept and store RT Plan from MLC Variable Aperture Arc Beam Producer |

### 3.27.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.27.4 Interaction Diagram

Archive

MLC Variable Aperture Arc Beam Producer

C-STORE (RT Plan)

#### 3.27.4.1 MLC Variable Aperture Arc Beam Storage

##### 3.27.4.1.1 Trigger Events

The MLC Variable Aperture Arc Beam Producer transfers the plan to the Archive once the plan is created and the dose calculation is finished.

##### 3.27.4.1.2 Message Semantics

The MLC Variable Aperture Arc Beam Producer uses the DICOM C-STORE message to transfer the plan.

The MLC Variable Aperture Arc Beam Producer is the DICOM Storage SCU and the Archive is the DICOM Storage SCP.

The MLC Variable Aperture Arc Beam Producer may create a new series containing the plan or may use an existing series, where previous plan(s) are contained.

The study where the series of the plan is contained shall be the same study as the one containing the structure set referenced in the plan.

###### 3.27.4.1.2.1 Storage of RT Plan containing a MLC Variable Aperture Arc Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.27.4.1.2.2 Optional Modifiers

The MLC Variable Aperture Arc Beam Producer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.27.4.1.3 Expected Actions

The Archive stores the RT Plan.

### 3.27.5 Security Considerations

There are no specific security considerations.

## 3.28 TPPC-10: MLC Variable Aperture Arc Beam Retrieval

### 3.28.1 Scope

In the MLC Variable Aperture Arc Beam Retrieval transaction, a consumer of an RT Plan that incorporates the beam technique identified in TPPC-09: MLC Variable Aperture Arc Beam Storage, retrieves the plan from the archive.

### 3.28.2 Use Case Roles

MLC Variable Aperture Arc Beam Retrieval

Archive

MLC Variable Aperture Arc Beam Consumer

| Actor: | MLC Variable Aperture Arc Beam Consumer |
| --- | --- |
| Role: | Stores plan transmitted from Archive |
| Actor: | Archive |
| Role: | Transmits Plan to MLC Variable Aperture Arc Beam Consumer |

### 3.28.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.28.4 Interaction Diagram

Archive

MLC Variable Aperture Arc Beam Consumer

C_STORE (RT Plan)

#### 3.28.4.1 MLC Variable Aperture Arc Beam Retrieval

##### 3.28.4.1.1 Trigger Events

The Archive transfers the plan to the MLC Variable Aperture Arc Beam Consumer.

##### 3.28.4.1.2 Message Semantics

The Archive uses the DICOM C-STORE message to transfer the plan.

The Archive is the DICOM Storage SCU and the MLC Variable Aperture Arc Beam Consumer is the DICOM Storage SCP.

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.28.4.1.2.1 Storage of RT Plan containing a MLC Variable Aperture Arc Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.28.4.1.2.2 Optional Modifiers

The MLC Variable Aperture Arc Beam Consumer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.28.4.1.3 Expected Actions

The MLC Variable Aperture Arc Beam Consumer stores the RT Plan.

### 3.28.5 Security Considerations

There are no specific security considerations.

## 3.29 TPPC-11 Hard Wedge Beam Storage

### 3.29.1 Scope

In the Hard Wedge Beam Storage transaction, a Producer of an RT Plan that incorporates the beam technique identified in TPPC-11: Hard Wedge Beam Storage stores the plan to the archive

### 3.29.2 Use Case Roles

Hard Wedge Beam Storage

Archive

Hard Wedge Beam Producer

| Actor: | Hard Wedge Beam Producer |
| --- | --- |
| Role: | Creates Hard Wedge Beam RT Plan and stores plan to an RT Archive |
| Actor: | Archive |
| Role: | Accept and store RT Plan from Hard Wedge Beam Producer |

### 3.29.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.29.4 Interaction Diagram

Archive

Hard Wedge Beam Producer

C-STORE (RT Plan)

#### 3.29.4.1 Hard Wedge Beam Storage

##### 3.29.4.1.1 Trigger Events

The Hard Wedge Beam Producer transfers the plan to the Archive once the plan is created and the dose calculation is finished.

##### 3.29.4.1.2 Message Semantics

The Hard Wedge Beam Producer uses the DICOM C-STORE message to transfer the plan.

The Hard Wedge Beam Producer is the DICOM Storage SCU and the Archive is the DICOM Storage SCP.

The Hard Wedge Beam Producer may create a new series containing the plan or may use an existing series, where previous plan(s) are contained.

The study where the series of the plan is contained shall be the same study as the one containing the structure set referenced in the plan.

###### 3.29.4.1.2.1 Storage of RT Plan containing a Hard Wedge Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.29.4.1.2.2 Optional Modifiers

The Hard Wedge Beam Producer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Compensator Beam Modifier | 7.4.4.3.3 |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.29.4.1.3 Expected Actions

The Archive stores the RT Plan.

### 3.29.5 Security Considerations

There are no specific security considerations.

## 3.30 TPPC-12: Hard Wedge Beam Retrieval

### 3.30.1 Scope

In the Hard Wedge Beam Retrieval transaction, a consumer of an RT Plan that incorporates the beam technique identified in TPPC-11: Hard Wedge Beam Storage, retrieves the plan from the archive.

### 3.30.2 Use Case Roles

Hard Wedge Beam Retrieval

Archive

Hard Wedge Beam Consumer

| Actor: | Hard Wedge Beam Consumer |
| --- | --- |
| Role: | Stores plan transmitted from Archive |
| Actor: | Archive |
| Role: | Transmits Plan to Hard Wedge Beam Consumer |

### 3.30.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.30.4 Interaction Diagram

Archive

Hard Wedge Beam Consumer

C_STORE (RT Plan)

#### 3.30.4.1 Hard Wedge Beam Retrieval

##### 3.30.4.1.1 Trigger Events

The Archive transfers the plan to the Hard Wedge Beam Consumer.

##### 3.30.4.1.2 Message Semantics

The Archive uses the DICOM C-STORE message to transfer the plan.

The Archive is the DICOM Storage SCU and the Hard Wedge Beam Consumer is the DICOM Storage SCP.

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.30.4.1.2.1 Storage of RT Plan containing a Hard Wedge Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.30.4.1.2.2 Optional Modifiers

The Hard Wedge Beam Consumer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Compensator Beam Modifier | 7.4.4.3.3 |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.30.4.1.3 Expected Actions

The Hard Wedge Beam Consumer stores the RT Plan.

### 3.30.5 Security Considerations

There are no specific security considerations.

## 3.31 TPPC-13 Virtual Wedge Beam Storage

### 3.31.1 Scope

In the Virtual Wedge Beam Storage transaction, a Producer of an RT Plan that incorporates the beam technique identified in TPPC-13: Virtual Wedge Beam Storage stores the plan to the archive

### 3.31.2 Use Case Roles

Virtual Wedge Beam Storage

Archive

Virtual Wedge Beam Producer

| Actor: | Virtual Wedge Beam Producer |
| --- | --- |
| Role: | Creates Virtual Wedge Beam RT Plan and stores plan to an RT Archive |
| Actor: | Archive |
| Role: | Accept and store RT Plan from Virtual Wedge Beam Producer |

### 3.31.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.31.4 Interaction Diagram

Archive

Virtual Wedge Beam Producer

C-STORE (RT Plan)

#### 3.31.4.1 Virtual Wedge Beam Storage

##### 3.31.4.1.1 Trigger Events

The Virtual Wedge Beam Producer transfers the plan to the Archive once the plan is created and the dose calculation is finished.

##### 3.31.4.1.2 Message Semantics

The Virtual Wedge Beam Producer uses the DICOM C-STORE message to transfer the plan.

The Virtual Wedge Beam Producer is the DICOM Storage SCU and the Archive is the DICOM Storage SCP.

The Virtual Wedge Beam Producer may create a new series containing the plan or may use an existing series, where previous plan(s) are contained.

The study where the series of the plan is contained shall be the same study as the one containing the structure set referenced in the plan.

###### 3.31.4.1.2.1 Storage of RT Plan containing a Virtual Wedge Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

A virtual wedge does not actually have an actual wedge angle (300A,00D5) in the same sense as a physical wedge. Most treatment planning systems, however, incorporate the concept of such an angle for reference during the planning and delivery process. This ‘physical’ setting of the delivery device should be placed in the wedge angle (300A,00D5) attribute.

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.31.4.1.2.2 Optional Modifiers

The Virtual Wedge Beam Producer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Hard Wedge Beam Modifier | 7.4.4.3.4 |
| Compensator Beam Modifier | 7.4.4.3.3 |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.31.4.1.3 Expected Actions

The Archive stores the RT Plan.

### 3.31.5 Security Considerations

There are no specific security considerations.

## 3.32 TPPC-14: Virtual Wedge Beam Retrieval

### 3.32.1 Scope

In the Virtual Wedge Beam Retrieval transaction, a consumer of an RT Plan that incorporates the beam technique identified in TPPC-13: Virtual Wedge Beam Storage, retrieves the plan from the archive.

### 3.32.2 Use Case Roles

Virtual Wedge Beam Retrieval

Archive

Virtual Wedge Beam Consumer

| Actor: | Virtual Wedge Beam Consumer |
| --- | --- |
| Role: | Stores plan transmitted from Archive |
| Actor: | Archive |
| Role: | Transmits Plan to Virtual Wedge Beam Consumer |

### 3.32.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.32.4 Interaction Diagram

Archive

Virtual Wedge Beam Consumer

C_STORE (RT Plan)

#### 3.32.4.1 Virtual Wedge Beam Retrieval

##### 3.32.4.1.1 Trigger Events

The Archive transfers the plan to the Virtual Wedge Beam Consumer.

##### 3.32.4.1.2 Message Semantics

The Archive uses the DICOM C-STORE message to transfer the plan.

The Archive is the DICOM Storage SCU and the Virtual Wedge Beam Consumer is the DICOM Storage SCP.

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.32.4.1.2.1 Storage of RT Plan containing a Virtual Wedge Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.32.4.1.2.2 Optional Modifiers

The Virtual Wedge Beam Consumer may support the following optional:

| Optional Modifiers | Section |
| --- | --- |
| Hard Wedge Beam Modifier | 7.4.4.3.4 |
| Compensator Beam Modifier | 7.4.4.3.3 |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.32.4.1.3 Expected Actions

The Virtual Wedge Beam Consumer stores the RT Plan.

### 3.32.5 Security Considerations

There are no specific security considerations.

## 3.33 TPPC-15 Motorized Wedge Beam Storage

### 3.33.1 Scope

In the Motorized Wedge Beam Storage transaction, a Producer of an RT Plan that incorporates the beam technique identified in TPPC-15: Virtual Wedge Beam Storage stores the plan to the archive

### 3.33.2 Use Case Roles

Motorized Wedge Beam Storage

Archive

Motorized Wedge Beam Producer

| Actor: | Motorized Wedge Beam Producer |
| --- | --- |
| Role: | Creates Motorized Wedge Beam RT Plan and stores plan to an RT Archive |
| Actor: | Archive |
| Role: | Accept and store RT Plan from Motorized Wedge Beam Producer |

### 3.33.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.33.4 Interaction Diagram

Archive

Motorized Wedge Beam Producer

C-STORE (RT Plan)

#### 3.33.4.1 Motorized Wedge Beam Storage

##### 3.33.4.1.1 Trigger Events

The Motorized Wedge Beam Producer transfers the plan to the Archive once the plan is created and the dose calculation is finished.

##### 3.33.4.1.2 Message Semantics

The Motorized Wedge Beam Producer uses the DICOM C-STORE message to transfer the plan.

The Motorized Wedge Beam Producer is the DICOM Storage SCU and the Archive is the DICOM Storage SCP.

The Motorized Wedge Beam Producer may create a new series containing the plan or may use an existing series, where previous plan(s) are contained.

The study where the series of the plan is contained shall be the same study as the one containing the structure set referenced in the plan.

###### 3.33.4.1.2.1 Storage of RT Plan containing a Motorized Wedge Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

In the delivery of a motorized wedge beam the wedge angle (300A,00D5) could represent either the effective angle of the total beam delivery or the angle of the physical wedge moved into the beam. For the TPPC Profile, the physical angle of the motorized wedge should be placed into the wedge angle (300A,00D5) attribute.

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.33.4.1.2.2 Optional Modifiers

The Motorized Wedge Beam Producer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Hard Wedge Beam Modifier | 7.4.4.3.4 |
| Compensator Beam Modifier | 7.4.4.3.3 |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.33.4.1.3 Expected Actions

The Archive stores the RT Plan.

### 3.33.5 Security Considerations

There are no specific security considerations.

## 3.34 TPPC-16: Motorized Wedge Beam Retrieval

### 3.34.1 Scope

In the Motorized Wedge Beam Retrieval transaction, a consumer of an RT Plan that incorporates the beam technique identified in TPPC-15: Motorized Wedge Beam Storage, retrieves the plan from the archive.

### 3.34.2 Use Case Roles

Motorized Wedge Beam Retrieval

Archive

Motorized Wedge Beam Consumer

| Actor: | Motorized Wedge Beam Consumer |
| --- | --- |
| Role: | Stores plan transmitted from Archive |
| Actor: | Archive |
| Role: | Transmits Plan to Motorized Wedge Beam Consumer |

### 3.34.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.34.4 Interaction Diagram

Archive

Motorized Wedge Beam Consumer

C_STORE (RT Plan)

#### 3.34.4.1 Motorized Wedge Beam Retrieval

##### 3.34.4.1.1 Trigger Events

The Archive transfers the plan to the Motorized Wedge Beam Consumer.

##### 3.34.4.1.2 Message Semantics

The Archive uses the DICOM C-STORE message to transfer the plan.

The Archive is the DICOM Storage SCU and the Motorized Wedge Beam Consumer is the DICOM Storage SCP.

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.34.4.1.2.1 Storage of RT Plan containing a Motorized Wedge Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.34.4.1.2.2 Optional Modifiers

The Motorized Wedge Beam Consumer may support the following optional:

| Optional Modifiers | Section |
| --- | --- |
| Hard Wedge Beam Modifier | 7.4.4.3.4 |
| Compensator Beam Modifier | 7.4.4.3.3 |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.34.4.1.3 Expected Actions

The Motorized Wedge Beam Consumer stores the RT Plan.

### 3.34.5 Security Considerations

There are no specific security considerations.

## 3.35 TPPC-17 Static Electron Beam Storage

### 3.35.1 Scope

In the Static Electron Beam Storage transaction, a Producer of an RT Plan that incorporates the beam technique identified in TPPC-17: Static Electron Beam Storage stores the plan to the archive

### 3.35.2 Use Case Roles

Static Electron Beam Storage

Archive

Static Electron Beam Producer

| Actor: | Static Electron Beam Producer |
| --- | --- |
| Role: | Creates Static Electron Beam RT Plan and stores plan to an RT Archive |
| Actor: | Archive |
| Role: | Accept and store RT Plan from Static Electron Beam Producer |

### 3.35.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.35.4 Interaction Diagram

Archive

Static Electron Beam Producer

C-STORE (RT Plan)

#### 3.35.4.1 Static Electron Beam Storage

##### 3.35.4.1.1 Trigger Events

The Static Electron Beam Producer transfers the plan to the Archive once the plan is created and the dose calculation is finished.

##### 3.35.4.1.2 Message Semantics

The Static Electron Beam Producer uses the DICOM C-STORE message to transfer the plan.

The Static Electron Beam Producer is the DICOM Storage SCU and the Archive is the DICOM Storage SCP.

The Static Electron Beam Producer may create a new series containing the plan or may use an existing series, where previous plan(s) are contained.

The study where the series of the plan is contained shall be the same study as the one containing the structure set referenced in the plan.

###### 3.35.4.1.2.1 Storage of RT Plan containing a Static Electron Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in Section 7.3.1.1.

###### 3.35.4.1.2.2 Optional Modifiers

The Static Electron Beam Producer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Compensator Beam Modifier | 7.4.4.3.3 |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.35.4.1.3 Expected Actions

The Archive stores the RT Plan.

### 3.35.5 Security Considerations

There are no specific security considerations.

## 3.36 TPPC-18: Static Electron Beam Retrieval

### 3.36.1 Scope

In the Static Electron Beam Retrieval transaction, a consumer of an RT Plan that incorporates the beam technique identified in TPPC-17: Static Electron Beam Storage, retrieves the plan from the archive.

### 3.36.2 Use Case Roles

Static Electron Beam Retrieval

Archive

Static Electron Beam Consumer

| Actor: | Static Electron Beam Consumer |
| --- | --- |
| Role: | Stores plan transmitted from Archive |
| Actor: | Archive |
| Role: | Transmits Plan to Static Electron Beam Consumer |

### 3.36.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.36.4 Interaction Diagram

Archive

Static Electron Beam Consumer

C_STORE (RT Plan)

#### 3.36.4.1 Static Electron Beam Retrieval

##### 3.36.4.1.1 Trigger Events

The Archive transfers the plan to the Static Electron Beam Consumer.

##### 3.36.4.1.2 Message Semantics

The Archive uses the DICOM C-STORE message to transfer the plan.

The Archive is the DICOM Storage SCU and the Static Electron Beam Consumer is the DICOM Storage SCP.

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.36.4.1.2.1 Storage of RT Plan containing a Static Electron Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.36.4.1.2.2 Optional Modifiers

The Static Electron Beam Consumer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Compensator Beam Modifier | 7.4.4.3.3 |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.36.4.1.3 Expected Actions

The Static Electron Beam Consumer stores the RT Plan.

### 3.36.5 Security Considerations

There are no specific security considerations.

## 3.37 TPPC-19 Step & Shoot Beam Storage

### 3.37.1 Scope

In the Step & Shoot Beam Storage transaction, a Producer of an RT Plan that incorporates the beam technique identified in TPPC-19: Step & Shoot Beam Storage stores the plan to the archive

### 3.37.2 Use Case Roles

Step & Shoot Beam Storage

Archive

Step & Shoot Beam Producer

| Actor: | Step & Shoot Beam Producer |
| --- | --- |
| Role: | Creates Step & Shoot Beam RT Plan and stores plan to an RT Archive |
| Actor: | Archive |
| Role: | Accept and store RT Plan from Step & Shoot Beam Producer |

### 3.37.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.37.4 Interaction Diagram

Archive

Step & Shoot Beam Producer

C-STORE (RT Plan)

#### 3.37.4.1 Step & Shoot Beam Storage

##### 3.37.4.1.1 Trigger Events

The Step & Shoot Beam Producer transfers the plan to the Archive once the plan is created and the dose calculation is finished.

##### 3.37.4.1.2 Message Semantics

The Step & Shoot Beam Producer uses the DICOM C-STORE message to transfer the plan.

The Step & Shoot Beam Producer is the DICOM Storage SCU and the Archive is the DICOM Storage SCP.

The Step & Shoot Beam Producer may create a new series containing the plan or may use an existing series, where previous plan(s) are contained.

The study where the series of the plan is contained shall be the same study as the one containing the structure set referenced in the plan.

###### 3.37.4.1.2.1 Storage of RT Plan containing a Step & Shoot Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.37.4.1.2.2 Optional Modifiers

The Step & Shoot Beam Producer may support the following optional :

| Optional Modifiers | Section |
| --- | --- |
| Hard Wedge Beam Modifier | 7.4.4.3.4 |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.37.4.1.3 Expected Actions

The Archive stores the RT Plan.

### 3.37.5 Security Considerations

There are no specific security considerations.

## 3.38 TPPC-20: Step & Shoot Beam Retrieval

### 3.38.1 Scope

In the Step & Shoot Beam Retrieval transaction, a consumer of an RT Plan that incorporates the beam technique identified in TPPC-19: Step & Shoot Beam Storage, retrieves the plan from the archive.

### 3.38.2 Use Case Roles

Step & Shoot Beam Retrieval

Archive

Step & Shoot Beam Consumer

| Actor: | Step & Shoot Beam Consumer |
| --- | --- |
| Role: | Stores plan transmitted from Archive |
| Actor: | Archive |
| Role: | Transmits Plan to Step & Shoot Beam Consumer |

### 3.38.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.38.4 Interaction Diagram

Archive

Step & Shoot Beam Consumer

C_STORE (RT Plan)

#### 3.38.4.1 Step & Shoot Beam Retrieval

##### 3.38.4.1.1 Trigger Events

The Archive transfers the plan to the Step & Shoot Beam Consumer.

##### 3.38.4.1.2 Message Semantics

The Archive uses the DICOM C-STORE message to transfer the plan.

The Archive is the DICOM Storage SCU and the Step & Shoot Beam Consumer is the DICOM Storage SCP.

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.38.4.1.2.1 Storage of RT Plan containing a Step & Shoot Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.38.4.1.2.2 Optional Modifiers

The Step & Shoot Beam Consumer may support the following optional:

| Optional Modifiers | Section |
| --- | --- |
| Hard Wedge Beam Modifier | 7.4.4.3.4 |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.38.4.1.3 Expected Actions

The Step & Shoot Beam Consumer stores the RT Plan.

### 3.38.5 Security Considerations

There are no specific security considerations.

## 3.39 TPPC-21 Sliding Window Beam Storage

### 3.39.1 Scope

In the Sliding Window Beam Storage transaction, a Producer of an RT Plan that incorporates the beam technique identified in TPPC-21: Sliding Window Beam Storage stores the plan to the archive

### 3.39.2 Use Case Roles

Sliding Window Beam Storage

Archive

Sliding Window Beam Producer

| Actor: | Sliding Window Beam Producer |
| --- | --- |
| Role: | Creates Sliding Window Beam RT Plan and stores plan to an RT Archive |
| Actor: | Archive |
| Role: | Accept and store RT Plan from Sliding Window Beam Producer |

### 3.39.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.39.4 Interaction Diagram

Archive

Sliding Window Beam Producer

C-STORE (RT Plan)

#### 3.39.4.1 Sliding Window Beam Storage

##### 3.39.4.1.1 Trigger Events

The Sliding Window Beam Producer transfers the plan to the Archive once the plan is created and the dose calculation is finished.

##### 3.39.4.1.2 Message Semantics

The Sliding Window Beam Producer uses the DICOM C-STORE message to transfer the plan.

The Sliding Window Beam Producer is the DICOM Storage SCU and the Archive is the DICOM Storage SCP.

The Sliding Window Beam Producer may create a new series containing the plan or may use an existing series, where previous plan(s) are contained.

The study where the series of the plan is contained shall be the same study as the one containing the structure set referenced in the plan.

###### 3.39.4.1.2.1 Storage of RT Plan containing a Sliding Window Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.39.4.1.2.2 Optional Modifiers

The Sliding Window Beam Producer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Hard Wedge Beam Modifier | 7.4.4.3.4 |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.39.4.1.3 Expected Actions

The Archive stores the RT Plan.

### 3.39.5 Security Considerations

There are no specific security considerations.

## 3.40 TPPC-22: Sliding Window Beam Retrieval

### 3.40.1 Scope

In the Sliding Window Beam Retrieval transaction, a consumer of an RT Plan that incorporates the beam technique identified in TPPC-21: Sliding Window Beam Storage, retrieves the plan from the archive.

### 3.40.2 Use Case Roles

Sliding Window Beam Retrieval

Archive

Sliding Window Beam Consumer

| Actor: | Sliding Window Beam Consumer |
| --- | --- |
| Role: | Stores plan transmitted from Archive |
| Actor: | Archive |
| Role: | Transmits Plan to Sliding Window Beam Consumer |

### 3.40.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.40.4 Interaction Diagram

Archive

Sliding Window Beam Consumer

C_STORE (RT Plan)

#### 3.40.4.1 Sliding Window Beam Retrieval

##### 3.40.4.1.1 Trigger Events

The Archive transfers the plan to the Sliding Window Beam Consumer.

##### 3.40.4.1.2 Message Semantics

The Archive uses the DICOM C-STORE message to transfer the plan.

The Archive is the DICOM Storage SCU and the Sliding Window Beam Consumer is the DICOM Storage SCP.

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.40.4.1.2.1 Storage of RT Plan containing a Sliding Window Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.40.4.1.2.2 Optional Modifiers

The Sliding Window Beam Consumer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Hard Wedge Beam Modifier | 7.4.4.3.4 |
| Bolus Beam Modifier | 7.4.4.3.1 |
| Block Beam Modifier | 7.4.4.3.2 |

##### 3.40.4.1.3 Expected Actions

The Sliding Window Beam Consumer stores the RT Plan.

### 3.40.5 Security Considerations

There are no specific security considerations.

## 3.41 TPPC-23 IMAT/VMAT Beam Storage

### 3.41.1 Scope

In the IMAT/VMAT Beam Storage transaction, a Producer of an RT Plan that incorporates the beam technique identified in TPPC-41: IMAT/VMAT Beam Storage stores the plan to the archive

### 3.41.2 Use Case Roles

IMAT/VMAT Beam Storage

Archive

IMAT/VMAT Beam Producer

| Actor: | IMAT/VMAT Beam Producer |
| --- | --- |
| Role: | Creates IMAT/VMAT Beam RT Plan and stores plan to an RT Archive |
| Actor: | Archive |
| Role: | Accept and store RT Plan from IMAT/VMAT Beam Producer |

### 3.41.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.41.4 Interaction Diagram

Archive

IMAT/VMAT Beam Producer

C-STORE (RT Plan)

#### 3.41.4.1 IMAT/VMAT Beam Storage

##### 3.41.4.1.1 Trigger Events

The IMAT/VMAT Beam Producer transfers the plan to the Archive once the plan is created and the dose calculation is finished.

##### 3.41.4.1.2 Message Semantics

The IMAT/VMAT Beam Producer uses the DICOM C-STORE message to transfer the plan.

The IMAT/VMAT Beam Producer is the DICOM Storage SCU and the Archive is the DICOM Storage SCP.

The IMAT/VMAT Beam Producer may create a new series containing the plan or may use an existing series, where previous plan(s) are contained.

The study where the series of the plan is contained shall be the same study as the one containing the structure set referenced in the plan.

###### 3.41.4.1.2.1 Storage of RT Plan containing a IMAT/VMAT Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.41.4.1.2.2 Optional Modifiers

The IMAT/VMAT Beam Producer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Bolus Beam Modifier | 7.4.4.3.1 |

##### 3.41.4.1.3 Expected Actions

The Archive stores the RT Plan.

### 3.41.5 Security Considerations

There are no specific security considerations.

## 3.42 TPPC-24: IMAT/VMAT Beam Retrieval

### 3.42.1 Scope

In the IMAT/VMAT Beam Retrieval transaction, a consumer of an RT Plan that incorporates the beam technique identified in TPPC-40: IMAT/VMAT Beam Storage, retrieves the plan from the archive.

### 3.42.2 Use Case Roles

IMAT/VMAT Beam Retrieval

Archive

IMAT/VMAT Beam Consumer

| Actor: | IMAT/VMAT Beam Consumer |
| --- | --- |
| Role: | Stores plan transmitted from Archive |
| Actor: | Archive |
| Role: | Transmits Plan to IMAT/VMAT Beam Consumer |

### 3.42.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.42.4 Interaction Diagram

Archive

IMAT/VMAT Beam Consumer

C_STORE (RT Plan)

#### 3.42.4.1 IMAT/VMAT Beam Retrieval

##### 3.42.4.1.1 Trigger Events

The Archive transfers the plan to the IMAT/VMAT Beam Consumer.

##### 3.42.4.1.2 Message Semantics

The Archive uses the DICOM C-STORE message to transfer the plan.

The Archive is the DICOM Storage SCU and the IMAT/VMAT Beam Consumer is the DICOM Storage SCP.

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.42.4.1.2.1 Storage of RT Plan containing a IMAT/VMAT Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.42.4.1.2.2 Optional Modifiers

The IMAT/VMAT Beam Consumer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Bolus Beam Modifier | 7.4.4.3.1 |

##### 3.42.4.1.3 Expected Actions

The IMAT/VMAT Beam Consumer stores the RT Plan.

### 3.42.5 Security Considerations

There are no specific security considerations.

## 3.43 TPPC-25 Photon Applicator Beam Storage

### 3.43.1 Scope

In the Photon Applicator Beam Storage transaction, a Producer of an RT Plan that incorporates the beam technique identified in TPPC-43: Photon Applicator Beam Storage stores the plan to the archive

### 3.43.2 Use Case Roles

Photon Applicator Beam Storage

Archive

Photon Applicator Beam Producer

| Actor: | Photon Applicator Beam Producer |
| --- | --- |
| Role: | Creates Photon Applicator Beam RT Plan and stores plan to an RT Archive |
| Actor: | Archive |
| Role: | Accept and store RT Plan from Photon Applicator Beam Producer |

### 3.43.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.43.4 Interaction Diagram

Archive

Photon Applicator Beam Producer

C-STORE (RT Plan)

#### 3.43.4.1 Photon Applicator Beam Storage

##### 3.43.4.1.1 Trigger Events

The Photon Applicator Beam Producer transfers the plan to the Archive once the plan is created and the dose calculation is finished.

##### 3.43.4.1.2 Message Semantics

The Photon Applicator Beam Producer uses the DICOM C-STORE message to transfer the plan.

The Photon Applicator Beam Producer is the DICOM Storage SCU and the Archive is the DICOM Storage SCP.

The Photon Applicator Beam Producer may create a new series containing the plan or may use an existing series, where previous plan(s) are contained.

The study where the series of the plan is contained shall be the same study as the one containing the structure set referenced in the plan.

###### 3.43.4.1.2.1 Storage of RT Plan containing a Photon Applicator Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.43.4.1.2.2 Optional Modifiers

The Photon Applicator Beam Producer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Bolus Beam Modifier | 7.4.4.3.1 |

##### 3.43.4.1.3 Expected Actions

The Archive stores the RT Plan.

### 3.43.5 Security Considerations

There are no specific security considerations.

## 3.44 TPPC-26: Photon Applicator Beam Retrieval

### 3.44.1 Scope

In the Photon Applicator Beam Retrieval transaction, a consumer of an RT Plan that incorporates the beam technique identified in TPPC-25: Photon Applicator Beam Storage, retrieves the plan from the archive.

### 3.44.2 Use Case Roles

Photon Applicator Beam Retrieval

Archive

Photon Applicator Beam Consumer

| Actor: | Photon Applicator Beam Consumer |
| --- | --- |
| Role: | Stores plan transmitted from Archive |
| Actor: | Archive |
| Role: | Transmits Plan to Photon Applicator Beam Consumer |

### 3.44.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.44.4 Interaction Diagram

Archive

Photon Applicator Beam Consumer

C_STORE (RT Plan)

#### 3.44.4.1 Photon Applicator Beam Retrieval

##### 3.44.4.1.1 Trigger Events

The Archive transfers the plan to the Photon Applicator Beam Consumer.

##### 3.44.4.1.2 Message Semantics

The Archive uses the DICOM C-STORE message to transfer the plan.

The Archive is the DICOM Storage SCU and the Photon Applicator Beam Consumer is the DICOM Storage SCP.

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.44.4.1.2.1 Storage of RT Plan containing a Photon Applicator Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.44.4.1.2.2 Optional Modifiers

The Photon Applicator Beam Consumer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Bolus Beam Modifier | 7.4.4.3.1 |

##### 3.44.4.1.3 Expected Actions

The Photon Applicator Beam Consumer stores the RT Plan.

### 3.44.5 Security Considerations

There are no specific security considerations.

## 3.45 TPPC-27 Photon Applicator Arc Beam Storage

### 3.45.1 Scope

In the Photon Applicator Arc Beam Storage transaction, a Producer of an RT Plan that incorporates the beam technique identified in TPPC-45: Photon Applicator Arc Beam Storage stores the plan to the archive

### 3.45.2 Use Case Roles

Photon Applicator Arc Beam Storage

Archive

Photon Applicator Arc Beam Producer

| Actor: | Photon Applicator Arc Beam Producer |
| --- | --- |
| Role: | Creates Basic Static Beam RT Plan and stores plan to an RT Archive |
| Actor: | Archive |
| Role: | Accept and store RT Plan from Photon Applicator Arc Beam Producer |

### 3.45.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.45.4 Interaction Diagram

Archive

Photon Applicator Arc Beam Producer

C-STORE (RT Plan)

#### 3.45.4.1 Photon Applicator Arc Beam Storage

##### 3.45.4.1.1 Trigger Events

The Photon Applicator Arc Beam Producer transfers the plan to the Archive once the plan is created and the dose calculation is finished.

##### 3.45.4.1.2 Message Semantics

The Photon Applicator Arc Beam Producer uses the DICOM C-STORE message to transfer the plan.

The Photon Applicator Arc Beam Producer is the DICOM Storage SCU and the Archive is the DICOM Storage SCP.

The Photon Applicator Arc Beam Producer may create a new series containing the plan or may use an existing series, where previous plan(s) are contained.

The study where the series of the plan is contained shall be the same study as the one containing the structure set referenced in the plan.

###### 3.45.4.1.2.1 Storage of RT Plan containing a Photon Applicator Arc Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.45.4.1.2.2 Optional Modifiers

The Photon Applicator Arc Beam Producer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Bolus Beam Modifier | 7.4.4.3.1 |

##### 3.45.4.1.3 Expected Actions

The Archive stores the RT Plan.

### 3.45.5 Security Considerations

There are no specific security considerations.

## 3.46 TPPC-28: Photon Applicator Arc Beam Retrieval

### 3.46.1 Scope

In the Photon Applicator Arc Beam Retrieval transaction, a consumer of an RT Plan that incorporates the beam technique identified in TPPC-27: Photon Applicator Arc Beam Storage, retrieves the plan from the archive.

### 3.46.2 Use Case Roles

Photon Applicator Arc Beam Retrieval

Archive

Photon Applicator Arc Beam Consumer

| Actor: | Photon Applicator Arc Beam Consumer |
| --- | --- |
| Role: | Stores plan transmitted from Archive |
| Actor: | Archive |
| Role: | Transmits Plan to Photon Applicator Arc Beam Consumer |

### 3.46.3 Referenced Standards

DICOM 2015a, PS 3.3: RT Modules, PS 3.4: Storage Service Class.

### 3.46.4 Interaction Diagram

Archive

Photon Applicator Arc Beam Consumer

C_STORE (RT Plan)

#### 3.46.4.1 Photon Applicator Arc Beam Retrieval

##### 3.46.4.1.1 Trigger Events

The Archive transfers the plan to the Photon Applicator Arc Beam Consumer.

##### 3.46.4.1.2 Message Semantics

The Archive uses the DICOM C-STORE message to transfer the plan.

The Archive is the DICOM Storage SCU and the Photon Applicator Arc Beam Consumer is the DICOM Storage SCP.

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.46.4.1.2.1 Storage of RT Plan containing a Photon Applicator Beam

Systems supporting the Treatment Planning - Plan Content Profile are required to support a number of attributes as described in the following tables and text. Many of these requirements build on attributes which are Type 2 or Type 3 in DICOM (such attributes are indicated with R+ or R+*).

All attributes in required modules for RT Plan as listed in chapters 7.3.2.1.

###### 3.46.4.1.2.2 Optional Modifiers

The Photon Applicator Arc Beam Consumer may support the following optional modifications:

| Optional Modifiers | Section |
| --- | --- |
| Bolus Beam Modifier | 7.4.4.3.1 |

##### 3.46.4.1.3 Expected Actions

The Photon Applicator Arc Beam Consumer stores the RT Plan.

### 3.46.5 Security Considerations

There are no specific security considerations.

Volume 3 – Content Modules

# 5 Namespaces and Vocabularies

No Namespaces and Vocabularies defined.

# 6 Content Modules

No Content Modules defined.

# DICOM Content Definition

## Conventions

<No change to framework>

## General Definitions

<No change to framework>

## IOD Definitions

This section defines each DICOM IOD used in the IHE Radiation Oncology domain in detail, specifying the standards used and the information defined.

### Prescription IODs

### Plan IODs

#### RT Plan IOD for Photon External Beam in Planning State

##### Referenced Standards

DICOM 2015a PS 3.3

##### IOD Definition

| IE | Module | Reference | Usage | IHE-RO Usage |
| --- | --- | --- | --- | --- |
| Patient | Patient | C.7.1.1 | M | M See 7.4.1.1.1 |
|  | Clinical Trial Subject | C.7.1.3 | U | U |
| Study | General Study | C.7.2.1 | M | M See 7.4.1.2.1 |
|  | Patient Study | C.7.2.2 | U | U |
|  | Clinical Trial Study | C.7.2.3 | U | U |
| Series | RT Series | C.8.8.1 | M | M See 7.4.1.4.1 |
|  | Clinical Trial Series | C.7.3.2 | U | U |
| Frame of Reference | Frame of Reference | C.7.4.1 | U - See Note. | R See7.4.1.7.1 |
| Equipment | General Equipment | C.7.5.1 | M | M See 7.4.1.5.1 |
| Plan | RT General Plan | C.8.8.9 | M | M See 7.4.3.1.1 |
|  | RT Prescription | C.8.8.10 | U | R See 7.4.3.2.1 |
|  | RT Tolerance Tables | C.8.8.11 | U | U |
|  | RT Patient Setup | C.8.8.12 | U | R See 7.4.5.3.1 |
|  | RT Fraction Scheme | C.8.8.13 | U | R See 7.4.3.3.1 and 7.4.3.3.2 |
|  | RT Beams | C.8.8.14 | C - Required if RT Fraction Scheme Module exists and Number of Beams (300A,0080) is greater than zero for one or more fraction groups | R Definitions see below |
|  | RT Brachy Application Setups | C.8.8.15 | C - Required if RT Fraction Scheme Module exists and Number of Brachy Application Setups (300A,00A0) is greater than zero for one or more fraction groups | Absent |
|  | Approval | C.8.8.16 | U | R |
|  | SOP Common | C.12.1 | M | M See 7.4.1.6.1 |

RT Beams Module is defined as follows:

| Beam Content Type | Section |
| --- | --- |
| Basic Static Beam | 7.4.4.1.1 |
| Basic Static MLC Beam | 7.4.4.1.2 |
| Arc Beam | 7.4.4.1.3 |
| MLC Fixed Aperture Arc Beam | 7.4.4.1.4 |
| MLC Variable Aperture Arc Beam | 7.4.4.1.5 |
| Hard Wedge Beam | 7.4.4.1.6 |
| Virtual Wedge Beam | 7.4.4.1.7 |
| Motorized Wedge Beam | 7.4.4.1.8 |
| Static Electron Beam | 7.4.4.1.9 |
| Step & Shoot Beam | 7.4.4.1.10 |
| Sliding Window Beam | 7.4.4.1.11 |
| IMAT/VMAT Beam | 7.4.4.1.12 |
| Photon Applicator Beam | 7.4.4.1.13 |
| Photon Applicator Arc Beam | 7.4.4.1.14 |

## Module Definitions

This section defines each DICOM Module used in the IHE Radiation Oncology domain in detail, specifying the standards used and the information defined.

When more than 1 type is shown for an attribute, the order is Producer / Consumer

### General Modules

#### Patient Module

##### Patient Module Base Content

<No change to framework>

#### Study Module

##### Study Module Base Content

<No change to framework>

#### General Series Module

##### General Series Module Base Content

<No change to framework>

#### RT Series Module

##### RT Series Module Base Content

<No change to framework>

#### Equipment Module

##### Equipment Module Base Content

<No change to framework>

#### SOP Common Module

##### SOP Common Module Base Content

<No change to framework>

#### Frame of Reference Module

##### Frame of Reference Module Base Content

<No change to framework>

### Workflow-Related Modules

### General Plan-Related Modules

#### General Plan Module

##### General Plan Module Base Content

<No change to framework>

#### RT Prescription Module

##### RT Prescription Module Base Content

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Module Definition

| Attribute | Tag | Type | Attribute Note |
| --- | --- | --- | --- |
| Dose Reference Sequence | (300A,0010) | R+* | See Dose reference requirements in the RT Beams module for the TPPC transactions. |
| > Dose Reference UID | (300A,0013) | R+* |  |
| > Dose Reference Description | (300A,0016) | R+ |  |

#### RT Fraction Scheme Module

This section is present only to convey the envisioned section numbering.

##### RT Fraction Scheme Module for Consistent Dose

This section is present only to convey the envisioned section numbering.

##### RT Fraction Scheme Module for Delivery

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Module Definition

| Attribute | Tag | Type | Attribute Note |
| --- | --- | --- | --- |
| Fraction Group Sequence | (300A,0070) | R+* | Shall have only a single item in the sequence |
| > Number of Fractions Planned | (300A,0078) | R+ |  |
| > Referenced Beam Sequence | (300C,0004) | R+* |  |
| >> Referenced Dose Reference UID | (300A,0083) | R+* | Identifies the Dose Reference specified by Dose Reference UID (300A,0013) in the Dose Reference Sequence (300A,0010) in the RT Prescription Module which specifies the primary target for the current Beam. If present shall have a value that is present in the Dose Reference Sequence. |
| >> Beam Dose | (300A,0084) | -/R+ R+/O+* | A TMS Actor is required to consume and process this value. A beam producer/consumer actor (e.g., a TPS) may consume this value and is required to produce it. |
| >> Beam Dose Specification Point | (300A,0082) | R+ |  |
| >> Beam Meterset | (300A,0086) | R+ |  |

### Plan-Related Modules in Planning

This section is present only to convey the envisioned section numbering.

#### Specific RT BeamType Specifications

##### RT Beams Module for Basic Static Beam

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Module Definition

| Attribute | Tag | Beam Technique |  |
| --- | --- | --- | --- |
|  |  | Basic Static |  |
|  |  | Presence | Specific Rules |
| Beam Sequence | (300A,00B0) | R+* |  |
| > Beam Number | (300A,00C0) | R+* | Shall be >= 1. |
| > Beam Name | (300A,00C2) | R+ |  |
| > Beam Type | (300A,00C4) | R+* | Shall be STATIC. |
| > Radiation Type | (300A,00C6) | R+* | Shall be PHOTON. |
| > High-Dose Technique Type | (300A,00C7) | O+* | If present, must be handled safely |
| > Primary Fluence Mode Sequence | (3002,0050) | R+* |  |
| >> Fluence Mode | (3002,0051) | D |  |
| >> Fluence Mode ID | (3002,0052) | D |  |
| > Treatment Machine Name | (300A,00B2) | R+* | Shall be constant. |
| > Primary Dosimeter Unit | (300A,00B3) | R+ | Shall be MU. |
| > Source-Axis Distance | (300A,00B4) | R+* |  |
| > Beam Limiting Device Sequence | (300A,00B6) | R+* |  |
| >> RT Beam Limiting Device Type | (300A,00B8) | R+* | Shall be 2 jaws, MLC shall not be present |
| >> Leaf Position Boundaries | (300A,00BE) | O+* | NA (no MLC) May or may not be present for jaws, may be ignored for jaws |
| > Referenced Patient Setup Number | (300C,006A) | R+* | Shall be >= 1. |
| > Treatment Delivery Type | (300A,00CE) | R+* |  |
| > Number of Wedges | (300A,00D0) | R+* | Shall be 0. |
| > Number of Compensators | (300A,00E0) | R+* | Shall be 0 or 1. If 1, see Compensator Beam Modifier. |
| > Number of Boli | (300A, 00ED) | R+* | Shall be >= 0. If > 0, see Bolus Beam Modifier. |
| > Number of Blocks | (300A,00F0) | R+* | Shall be 0-8. If > 0, see Block Beam Modifier. |
| > Applicator Sequence | (300A,0107) | O+ | Shall not be present. |
| > Final Cumulative Meterset Weight | (300A,010E) | R+* |  |
| > Number of Control Points | (300A,0110) | R+* | Shall be 2. |
| > Control Point Sequence | (300A,0111) | R+* |  |
| >> Cumulative Meterset Weight | (300A,0134) | R+ |  |
| >> Referenced Dose Reference Sequence | (300C,0050) | -/R+* R+/O+* | A TMS Actor is required to consume and process this value. A beam producer/consumer actor (e.g., a TPS) may consume this value and is required to produce it. Shall have at least one item for target dose accumulation. |
| >>> Cumulative Dose Reference Coefficient | (300A,010C) | R+* | Shall be present. |
| >> Nominal Beam Energy | (300A,0114) | R+ | Shall be constant. |
| >> Dose Rate Set | (300A,0115) | R+ | Shall be constant. |
| >> Wedge Position Sequence | (300A,0116) | R+* | Shall not be present |
| >> Beam Limiting Device Position Sequence | (300A,011A) | R+* | Shall be consistent with the Beam Limiting Device Sequence (300A,00B6). |
| >>>Leaf/Jaw Positions | (300A,011C) | R+* |  |
| >> Gantry Angle | (300A,011E) | R+* | Shall be constant. |
| >> Gantry Rotation Direction | (300A,011F) | R+* | Shall be NONE. |
| >> Gantry Pitch Angle | (300A,014A) | O+* | If not present, shall be assumed to be in the zero position. If present, shall be zero. |
| >> Gantry Pitch Rotation Direction | (300A,014C) | O+* | If present, shall be NONE. |
| >> Beam Limiting Device Angle | (300A,0120) | R+* | Shall be constant. |
| >> Beam Limiting Device Rotation Direction | (300A,0121) | R+* | Shall be NONE. |
| < Insert Control Point Sequence Fixed Attributes List > (See 7.4.4.2.1) |  |  |  |
| >> Isocenter Position | (300A,012C) | R+ | Shall be constant for all CPs. |

##### RT Beams Module for Basic Static MLC Beam

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Module Definition

| Attribute | Tag | Beam Technique |  |
| --- | --- | --- | --- |
|  |  | Basic Static MLC |  |
|  |  | Presence | Specific Rules |
| Beam Sequence | (300A,00B0) | R+* |  |
| > Beam Number | (300A,00C0) | R+* | Shall be >= 1. |
| > Beam Name | (300A,00C2) | R+ |  |
| > Beam Type | (300A,00C4) | R+* | Shall be STATIC. |
| > Radiation Type | (300A,00C6) | R+* | Shall be PHOTON. |
| > High-Dose Technique Type | (300A,00C7) | O+* | If present, must be handled safely |
| > Primary Fluence Mode Sequence | (3002,0050) | R+* |  |
| >> Fluence Mode | (3002,0051) | D |  |
| >> Fluence Mode ID | (3002,0052) | D |  |
| > Treatment Machine Name | (300A,00B2) | R+* | Shall be constant. |
| > Primary Dosimeter Unit | (300A,00B3) | R+ | Shall be MU. |
| > Source-Axis Distance | (300A,00B4) | R+* |  |
| > Beam Limiting Device Sequence | (300A,00B6) | R+* |  |
| >> RT Beam Limiting Device Type | (300A,00B8) | R+* | Shall have at least 1 MLC |
| >> Leaf Position Boundaries | (300A,00BE) | R+* | Shall be present for MLCs May or may not be present for jaws, may be ignored for jaws |
| > Referenced Patient Setup Number | (300C,006A) | R+* | Shall be >= 1. |
| > Treatment Delivery Type | (300A,00CE) | R+* |  |
| > Number of Wedges | (300A,00D0) | R+* | Shall be 0. |
| > Number of Compensators | (300A,00E0) | R+* | Shall be 0 or 1. If 1, see Compensator Beam Modifier. |
| > Number of Boli | (300A, 00ED) | R+* | Shall be >= 0. If > 0, see Bolus Beam Modifier. |
| > Number of Blocks | (300A,00F0) | R+* | Shall be 0-8. If > 0, see Block Beam Modifier. |
| > Applicator Sequence | (300A,0107) | R+* | Shall not be present. |
| > Final Cumulative Meterset Weight | (300A,010E) | R+* |  |
| > Number of Control Points | (300A,0110) | R+* | Shall be 2. |
| > Control Point Sequence | (300A,0111) | R+* |  |
| >> Cumulative Meterset Weight | (300A,0134) | R+ |  |
| >> Referenced Dose Reference Sequence | (300C,0050) | -/R+* R+/O+* | A TMS Actor is required to consume and process this value. A beam producer/consumer actor (e.g., a TPS) may consume this value and is required to produce it. Shall have at least one item for target dose accumulation. |
| >>> Cumulative Dose Reference Coefficient | (300A,010C) | R+* | Shall be present. |
| >> Nominal Beam Energy | (300A,0114) | R+ | Shall be constant. |
| >> Dose Rate Set | (300A,0115) | R+ | Shall be constant. |
| >> Wedge Position Sequence | (300A,0116) | R+* | Shall not be present |
| >> Beam Limiting Device Position Sequence | (300A,011A) | R+* | Shall be consistent with the Beam Limiting Device Sequence (300A,00B6). |
| >>>Leaf/Jaw Positions | (300A,011C) | R+* |  |
| >> Gantry Angle | (300A,011E) | R+* | Shall be constant. |
| >> Gantry Rotation Direction | (300A,011F) | R+* | Shall be NONE. |
| >> Gantry Pitch Angle | (300A,014A) | O+* | If not present, shall be assumed to be in the zero position. If present, shall be zero. |
| >> Gantry Pitch Rotation Direction | (300A,014C) | O+* | If present, shall be NONE. |
| >> Beam Limiting Device Angle | (300A,0120) | R+* | Shall be constant. |
| >> Beam Limiting Device Rotation Direction | (300A,0121) | R+* | Shall be NONE. |
| < Insert Control Point Sequence Fixed Attributes List > (See 7.4.4.2.1) |  |  |  |
| >> Isocenter Position | (300A,012C) | R+ | Shall be constant for all CPs. |

##### RT Beams Module for Arc Beam

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Module Definition

| Attribute | Tag | Beam Technique |  |
| --- | --- | --- | --- |
|  |  | Arc |  |
|  |  | Presence | Specific Rules |
| Beam Sequence | (300A,00B0) | R+* |  |
| > Beam Number | (300A,00C0) | R+* | Shall be >= 1. |
| > Beam Name | (300A,00C2) | R+ |  |
| > Beam Type | (300A,00C4) | R+* | Shall be DYNAMIC. |
| > Radiation Type | (300A,00C6) | R+* | Shall be PHOTON. |
| > High-Dose Technique Type | (300A,00C7) | O+* | If present, must be handled safely |
| > Primary Fluence Mode Sequence | (3002,0050) | R+* |  |
| >> Fluence Mode | (3002,0051) | D |  |
| >> Fluence Mode ID | (3002,0052) | D |  |
| > Treatment Machine Name | (300A,00B2) | R+* | Shall be constant. |
| > Primary Dosimeter Unit | (300A,00B3) | R+ | Shall be MU. |
| > Source-Axis Distance | (300A,00B4) | R+* |  |
| > Beam Limiting Device Sequence | (300A,00B6) | R+* |  |
| >> RT Beam Limiting Device Type | (300A,00B8) | R+* | Shall be 2 jaws, MLC shall not be present |
| >> Leaf Position Boundaries | (300A,00BE) | O+* | NA (no MLC) May or may not be present for jaws, may be ignored for jaws |
| > Referenced Patient Setup Number | (300C,006A) | R+* | Shall be >= 1. |
| > Treatment Delivery Type | (300A,00CE) | R+* |  |
| > Number of Wedges | (300A,00D0) | R+* | Shall be 0. |
| > Number of Compensators | (300A,00E0) | R+* | Shall be 0. |
| > Number of Boli | (300A, 00ED) | R+* | Shall be >= 0. If > 0, see Bolus Beam Modifier. |
| > Number of Blocks | (300A,00F0) | R+* | Shall be 0-8. If > 0, see Block Beam Modifier. |
| > Applicator Sequence | (300A,0107) | R+* | Shall not be present. |
| > Final Cumulative Meterset Weight | (300A,010E) | R+* |  |
| > Number of Control Points | (300A,0110) | R+* | Shall be 2. Skip arcs are not tested in this transaction. |
| > Control Point Sequence | (300A,0111) | R+* |  |
| >> Cumulative Meterset Weight | (300A,0134) | R+ |  |
| >> Referenced Dose Reference Sequence | (300C,0050) | -/R+* R+/O+* | A TMS Actor is required to consume and process this value. A beam producer/consumer actor (e.g., a TPS) may consume this value and is required to produce it Shall have at least one item for target dose accumulation. |
| >>> Cumulative Dose Reference Coefficient | (300A,010C) | R+* | Shall be present. |
| >> Nominal Beam Energy | (300A,0114) | R+ | Shall be constant. |
| >> Dose Rate Set | (300A,0115) | R+ | Shall be constant. |
| >> Wedge Position Sequence | (300A,0116) | R+* | Shall not be present |
| >> Beam Limiting Device Position Sequence | (300A,011A) | R+* | Shall be consistent with the Beam Limiting Device Sequence (300A,00B6). |
| >>>Leaf/Jaw Positions | (300A,011C) | R+* |  |
| >> Gantry Angle | (300A,011E) | R+* |  |
| >> Gantry Rotation Direction | (300A,011F) | R+* | Shall be CW or CC for Control Point 0. Can be NONE for Control Point 1. |
| >> Gantry Pitch Angle | (300A,014A) | O+* | If not present, shall be assumed to be in the zero position. If present, shall be zero. |
| >> Gantry Pitch Rotation Direction | (300A,014C) | O+* | If present, shall be NONE. |
| >> Beam Limiting Device Angle | (300A,0120) | R+* | Shall be constant. |
| >> Beam Limiting Device Rotation Direction | (300A,0121) | R+* | Shall be NONE. |
| < Insert Control Point Sequence Fixed Attributes List > (See 7.4.4.2.1) |  |  |  |
| >> Isocenter Position | (300A,012C) | R+ | Shall be constant for all CPs. |

##### RT Beams Module for MLC Fixed Aperture Arc Beam

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Module Definition

| Attribute | Tag | Beam Technique |  |
| --- | --- | --- | --- |
|  |  | MLC Fixed Aperture Arc |  |
|  |  | Presence | Specific Rules |
| Beam Sequence | (300A,00B0) | R+* |  |
| > Beam Number | (300A,00C0) | R+* | Shall be >= 1. |
| > Beam Name | (300A,00C2) | R+ |  |
| > Beam Type | (300A,00C4) | R+* | Shall be DYNAMIC. |
| > Radiation Type | (300A,00C6) | R+* | Shall be PHOTON. |
| > High-Dose Technique Type | (300A,00C7) | O+* | If present, must be handled safely |
| > Primary Fluence Mode Sequence | (3002,0050) | R+* |  |
| >> Fluence Mode | (3002,0051) | D |  |
| >> Fluence Mode ID | (3002,0052) | D |  |
| > Treatment Machine Name | (300A,00B2) | R+* | Shall be constant. |
| > Primary Dosimeter Unit | (300A,00B3) | R+ | Shall be MU. |
| > Source-Axis Distance | (300A,00B4) | R+* |  |
| > Beam Limiting Device Sequence | (300A,00B6) | R+* |  |
| >> RT Beam Limiting Device Type | (300A,00B8) | R+* | Shall have at least 1 MLC. |
| >> Leaf Position Boundaries | (300A,00BE) | R+* | Shall be present for MLCs. May or may not be present for jaws, may be ignored for jaws |
| > Referenced Patient Setup Number | (300C,006A) | R+* | Shall be >= 1. |
| > Treatment Delivery Type | (300A,00CE) | R+* |  |
| > Number of Wedges | (300A,00D0) | R+* | Shall be 0. |
| > Number of Compensators | (300A,00E0) | R+* | Shall be 0. |
| > Number of Boli | (300A, 00ED) | R+* | Shall be >= 0. If > 0, see Bolus Beam Modifier. |
| > Number of Blocks | (300A,00F0) | R+* | Shall be 0. |
| > Applicator Sequence | (300A,0107) | R+* | Shall not be present. |
| > Final Cumulative Meterset Weight | (300A,010E) | R+* |  |
| > Number of Control Points | (300A,0110) | R+* | Shall be 2. Skip arcs are not tested in this transaction. |
| > Control Point Sequence | (300A,0111) | R+* |  |
| >> Cumulative Meterset Weight | (300A,0134) | R+ |  |
| >> Referenced Dose Reference Sequence | (300C,0050) | -/R+* R+/O+* | A TMS Actor is required to consume and process this value. A beam producer/consumer actor (e.g., a TPS) may consume this value and is required to produce it Shall have at least one item for target dose accumulation. |
| >>> Cumulative Dose Reference Coefficient | (300A,010C) | R+* | Shall be present. |
| >> Nominal Beam Energy | (300A,0114) | R+ | Shall be constant. |
| >> Dose Rate Set | (300A,0115) | R+ | Shall be constant. |
| >> Wedge Position Sequence | (300A,0116) | R+* | Shall not be present |
| >> Beam Limiting Device Position Sequence | (300A,011A) | R+* | Shall be consistent with the Beam Limiting Device Sequence (300A,00B6). |
| >>>Leaf/Jaw Positions | (300A,011C) | R+* |  |
| >> Gantry Angle | (300A,011E) | R+* |  |
| >> Gantry Rotation Direction | (300A,011F) | R+* | Shall be CW or CC for Control Point 0. Can be NONE for Control Point 1. |
| >> Gantry Pitch Angle | (300A,014A) | O+* | If not present, shall be assumed to be in the zero position. If present, shall be zero. |
| >> Gantry Pitch Rotation Direction | (300A,014C) | O+* | If present, shall be NONE. |
| >> Beam Limiting Device Angle | (300A,0120) | R+* | Shall be constant. |
| >> Beam Limiting Device Rotation Direction | (300A,0121) | R+* | Shall be NONE. |
| < Insert Control Point Sequence Fixed Attributes List > (See 7.4.4.2.1) |  |  |  |
| >> Isocenter Position | (300A,012C) | R+ | Shall be constant for all CPs. |

##### RT Beams Module for MLC Variable Aperture Arc Beam

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Module Definition

| Attribute | Tag | Beam Technique |  |
| --- | --- | --- | --- |
|  |  | MLC Variable Aperture Arc |  |
|  |  | Presence | Specific Rules |
| Beam Sequence | (300A,00B0) | R+* |  |
| > Beam Number | (300A,00C0) | R+* | Shall be >= 1. |
| > Beam Name | (300A,00C2) | R+ |  |
| > Beam Type | (300A,00C4) | R+* | Shall be DYNAMIC. |
| > Radiation Type | (300A,00C6) | R+* | Shall be PHOTON. |
| > High-Dose Technique Type | (300A,00C7) | O+* | If present, must be handled safely |
| > Primary Fluence Mode Sequence | (3002,0050) | R+* |  |
| >> Fluence Mode | (3002,0051) | D |  |
| >> Fluence Mode ID | (3002,0052) | D |  |
| > Treatment Machine Name | (300A,00B2) | R+* | Shall be constant. |
| > Primary Dosimeter Unit | (300A,00B3) | R+ | Shall be MU. |
| > Source-Axis Distance | (300A,00B4) | R+* |  |
| > Beam Limiting Device Sequence | (300A,00B6) | R+* |  |
| >> RT Beam Limiting Device Type | (300A,00B8) | R+* | Shall be 2 jaws, or at least 1 jaw and 1 MLC. |
| >> Leaf Position Boundaries | (300A,00BE) | R+* | Shall be present for MLCs. May or may not be present for jaws, may be ignored for jaws |
| > Referenced Patient Setup Number | (300C,006A) | R+* | Shall be >= 1. |
| > Treatment Delivery Type | (300A,00CE) | R+* |  |
| > Number of Wedges | (300A,00D0) | R+* | Shall be 0. |
| > Number of Compensators | (300A,00E0) | R+* | Shall be 0. |
| > Number of Boli | (300A, 00ED) | R+* | Shall be >= 0. If > 0, see Bolus Beam Modifier. |
| > Number of Blocks | (300A,00F0) | R+* | Shall be 0-8. If > 0, see Block Beam Modifier. |
| > Applicator Sequence | (300A,0107) | R+* | Shall not be present. |
| > Final Cumulative Meterset Weight | (300A,010E) | R+* |  |
| > Number of Control Points | (300A,0110) | R+* | If the Consumer has a limit, it shall document this and safely handle input that exceeds that limit. |
| > Control Point Sequence | (300A,0111) | R+* |  |
| >> Cumulative Meterset Weight | (300A,0134) | R+ |  |
| >> Referenced Dose Reference Sequence | (300C,0050) | -/R+* R+/O+* | A TMS Actor is required to consume and process this value. A beam producer/consumer actor (e.g., a TPS) may consume this value and is required to produce it Shall have at least one item for target dose accumulation. |
| >>> Cumulative Dose Reference Coefficient | (300A,010C) | R+* | Shall be present. |
| >> Nominal Beam Energy | (300A,0114) | R+ | Shall be constant. |
| >> Dose Rate Set | (300A,0115) | R+ | Shall be constant. |
| >> Wedge Position Sequence | (300A,0116) | R+* | Shall not be present |
| >> Beam Limiting Device Position Sequence | (300A,011A) | R+* | Shall be consistent with the Beam Limiting Device Sequence (300A,00B6). |
| >>>Leaf/Jaw Positions | (300A,011C) | R+* |  |
| >> Gantry Angle | (300A,011E) | R+* |  |
| >> Gantry Rotation Direction | (300A,011F) | R+* | Shall be constant (CW or CC) for all CP except last one. Can be NONE for final CP |
| >> Gantry Pitch Angle | (300A,014A) | O+* | If not present, shall be assumed to be in the zero position. If present, shall be zero. |
| >> Gantry Pitch Rotation Direction | (300A,014C) | O+* | If present, shall be NONE. |
| >> Beam Limiting Device Angle | (300A,0120) | R+* | Shall be constant. |
| >> Beam Limiting Device Rotation Direction | (300A,0121) | R+* | Shall be NONE. |
| < Insert Control Point Sequence Fixed Attributes List > (See 7.4.4.2.1) |  |  |  |
| >> Isocenter Position | (300A,012C) | R+ | Shall be constant for all CPs. |

##### RT Beams Module for Hard Wedge Beam

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Module Definition

| Attribute | Tag | Beam Technique |  |
| --- | --- | --- | --- |
|  |  | Hard Wedge |  |
|  |  | Presence | Specific Rules |
| Beam Sequence | (300A,00B0) | R+* |  |
| > Beam Number | (300A,00C0) | R+* | Shall be >= 1. |
| > Beam Name | (300A,00C2) | R+ |  |
| > Beam Type | (300A,00C4) | R+* | Shall be STATIC. |
| > Radiation Type | (300A,00C6) | R+* | Shall be PHOTON. |
| > High-Dose Technique Type | (300A,00C7) | O+* | If present, must be handled safely |
| > Primary Fluence Mode Sequence | (3002,0050) | R+* |  |
| >> Fluence Mode | (3002,0051) | D |  |
| >> Fluence Mode ID | (3002,0052) | D |  |
| > Treatment Machine Name | (300A,00B2) | R+* | Shall be constant. |
| > Primary Dosimeter Unit | (300A,00B3) | R+ | Shall be MU. |
| > Source-Axis Distance | (300A,00B4) | R+* |  |
| > Beam Limiting Device Sequence | (300A,00B6) | R+* |  |
| >> RT Beam Limiting Device Type | (300A,00B8) | R+* | Shall have at least 2 jaws or at least 1 jaw and 1 MLC. |
| >> Leaf Position Boundaries | (300A,00BE) | R+* | Shall be present for MLCs,. May or may not be present for jaws, may be ignored for jaws |
| > Referenced Patient Setup Number | (300C,006A) | R+* | Shall be >= 1. |
| > Treatment Delivery Type | (300A,00CE) | R+* |  |
| > Number of Wedges | (300A,00D0) | R+* | Shall be 1. |
| >> Wedge Sequence | (300A,00D1) | R+* | Shall be present. |
| >> Wedge Type | (300A,00D3) | R+* | Shall be STANDARD (static) |
| >> Wedge ID | (300A,00D4) | R+ |  |
| >> Wedge Angle | (300A,00D5) | R+ |  |
| >> Wedge Orientation | (300A,00D8) | R+ |  |
| >> Source to Wedge Tray Distance | (300A,00DA) | R+ |  |
| > Number of Compensators | (300A,00E0) | R+* | Shall be 0 or 1. If 1, see Compensator Beam Modifier. |
| > Number of Boli | (300A, 00ED) | R+* | Shall be >= 0. If > 0, see Bolus Beam Modifier. |
| > Number of Blocks | (300A,00F0) | R+* | Shall be 0-8. If > 0, see Block Beam Modifier. |
| > Applicator Sequence | (300A,0107) | R+* | Shall not be present. |
| > Final Cumulative Meterset Weight | (300A,010E) | R+* |  |
| > Number of Control Points | (300A,0110) | R+* | Shall be 2. |
| > Control Point Sequence | (300A,0111) | R+* |  |
| >> Cumulative Meterset Weight | (300A,0134) | R+ |  |
| >> Referenced Dose Reference Sequence | (300C,0050) | -/R+* R+/O+* | A TMS Actor is required to consume and process this value. A beam producer/consumer actor (e.g., a TPS) may consume this value and is required to produce it Shall have at least one item for target dose accumulation. |
| >>> Cumulative Dose Reference Coefficient | (300A,010C) | R+* | Shall be present. |
| >> Nominal Beam Energy | (300A,0114) | R+ | Shall be constant. |
| >> Dose Rate Set | (300A,0115) | R+ | Shall be constant. |
| >> Wedge Position Sequence | (300A,0116) | R+* | Shall be present and consistent with the Wedge Sequence (300A,00D1). |
| >>> Wedge Position | (300A,0118) | R+* | Shall be IN. |
| >> Beam Limiting Device Position Sequence | (300A,011A) | R+* | Shall be consistent with the Beam Limiting Device Sequence (300A,00B6). |
| >>>Leaf/Jaw Positions | (300A,011C) | R+* |  |
| >> Gantry Angle | (300A,011E) | R+* | Shall be constant. |
| >> Gantry Rotation Direction | (300A,011F) | R+* | Shall be NONE. |
| >> Gantry Pitch Angle | (300A,014A) | O+* | If not present, shall be assumed to be in the zero position. If present, shall be zero. |
| >> Gantry Pitch Rotation Direction | (300A,014C) | O+* | If present, shall be NONE. |
| >> Beam Limiting Device Angle | (300A,0120) | R+* | Shall be constant. |
| >> Beam Limiting Device Rotation Direction | (300A,0121) | R+* | Shall be NONE. |
| < Insert Control Point Sequence Fixed Attributes List >(See 7.4.4.2.1) |  |  |  |
| >> Isocenter Position | (300A,012C) | R+ | Shall be constant for all CPs. |

##### RT Beams Module for Virtual Wedge Beam

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Module Definition

| Attribute | Tag | Beam Technique |  |
| --- | --- | --- | --- |
|  |  | Virtual Wedge |  |
|  |  | Presence | Specific Rules |
| Beam Sequence | (300A,00B0) | R+* |  |
| > Beam Number | (300A,00C0) | R+* | Shall be >= 1. |
| > Beam Name | (300A,00C2) | R+ |  |
| > Beam Type | (300A,00C4) | R+* | Shall be STATIC. |
| > Radiation Type | (300A,00C6) | R+* | Shall be PHOTON. |
| > High-Dose Technique Type | (300A,00C7) | O+* | If present, must be handled safely |
| > Primary Fluence Mode Sequence | (3002,0050) | R+* |  |
| >> Fluence Mode | (3002,0051) | D |  |
| >> Fluence Mode ID | (3002,0052) | D |  |
| > Treatment Machine Name | (300A,00B2) | R+* | Shall be constant. |
| > Primary Dosimeter Unit | (300A,00B3) | R+ | Shall be MU. |
| > Source-Axis Distance | (300A,00B4) | R+* |  |
| > Beam Limiting Device Sequence | (300A,00B6) | R+* |  |
| >> RT Beam Limiting Device Type | (300A,00B8) | R+* | Shall have at least 2 jaws or at least 1 jaw and 1 MLC. |
| >> Leaf Position Boundaries | (300A,00BE) | R+* | Shall be present for MLCs,. May or may not be present for jaws, may be ignored for jaws |
| > Referenced Patient Setup Number | (300C,006A) | R+* | Shall be >= 1. |
| > Treatment Delivery Type | (300A,00CE) | R+* |  |
| > Number of Wedges | (300A,00D0) | R+* | Shall be 1 or 2. If 2, see Hard Wedge Beam Modifier. |
| >>Wedge Sequence | (300A,00D1) | R+* | Shall be present if number of wedges is non-zero |
| >> Wedge Type | (300A,00D3) | R+* | Shall be DYNAMIC. Optional Hard Wedge shall be STANDARD |
| >> Wedge ID | (300A,00D4) | R+ |  |
| >> Wedge Angle | (300A,00D5) | RC+ | Shall be present if Wedge Type (300A,00D3) is STANDARD. May be present otherwise. |
| >>Effective Wedge Angle | (300A,00DE) | -/RC+ RC+/O+* | Shall be present if Wedge Type (300A,00D3) is DYNAMIC. A TMS Actor is required to consume and process this value. A beam producer/consumer actor (e.g., a TPS) may consume this value and is required to produce it |
| >> Wedge Orientation | (300A,00D8) | R+ |  |
| >> Source to Wedge Tray Distance | (300A,00DA) | RC+ | Shall be present if Wedge Type (300A,00D3) is STANDARD. |
| > Number of Compensators | (300A,00E0) | R+* | Shall be 0 or 1. If 1, see Compensator Beam Modifier. |
| > Number of Boli | (300A, 00ED) | R+* | Shall be >= 0. If > 0, see Bolus Beam Modifier. |
| > Number of Blocks | (300A,00F0) | R+* | Shall be 0-8. If > 0, see Block Beam Modifier. |
| > Applicator Sequence | (300A,0107) | R+* | Shall not be present. |
| > Final Cumulative Meterset Weight | (300A,010E) | R+* |  |
| > Number of Control Points | (300A,0110) | R+* | Shall be 2. |
| > Control Point Sequence | (300A,0111) | R+* |  |
| >> Cumulative Meterset Weight | (300A,0134) | R+ |  |
| >> Referenced Dose Reference Sequence | (300C,0050) | -/R+* R+/O+* | A TMS Actor is required to consume and process this value. A beam producer/consumer actor (e.g., a TPS) may consume this value and is required to produce it. Shall have at least one item for target dose accumulation. |
| >>> Cumulative Dose Reference Coefficient | (300A,010C) | R+* | Shall be present. |
| >> Nominal Beam Energy | (300A,0114) | R+ | Shall be constant. |
| >> Dose Rate Set | (300A,0115) | R+ | Shall be constant. |
| >> Wedge Position Sequence | (300A,0116) | R+* | Shall be present and consistent with the Wedge Sequence (300A,00D1). |
| >>> Wedge Position | (300A,0118) | R+* | Shall be IN. |
| >> Beam Limiting Device Position Sequence | (300A,011A) | R+* | Shall be consistent with the Beam Limiting Device Sequence (300A,00B6). |
| >>>Leaf/Jaw Positions | (300A,011C) | R+* |  |
| >> Gantry Angle | (300A,011E) | R+* | Shall be constant. |
| >> Gantry Rotation Direction | (300A,011F) | R+* | Shall be NONE. |
| >> Gantry Pitch Angle | (300A,014A) | O+* | If not present, shall be assumed to be in the zero position. If present, shall be zero. |
| >> Gantry Pitch Rotation Direction | (300A,014C) | O+* | If present, shall be NONE. |
| >> Beam Limiting Device Angle | (300A,0120) | R+* | Shall be constant. |
| >> Beam Limiting Device Rotation Direction | (300A,0121) | R+* | Shall be NONE. |
| < Insert Control Point Sequence Fixed Attributes List > (See 7.4.4.2.1) |  |  |  |
| >> Isocenter Position | (300A,012C) | R+ | Shall be constant for all CPs. |

##### RT Beams Module for Motorized Wedge Beam

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Module Definition

| Attribute | Tag | Beam Technique |  |
| --- | --- | --- | --- |
|  |  | Motorized Wedge |  |
|  |  | Presence | Specific Rules |
| Beam Sequence | (300A,00B0) | R+* |  |
| > Beam Number | (300A,00C0) | R+* | Shall be >= 1. |
| > Beam Name | (300A,00C2) | R+ |  |
| > Beam Type | (300A,00C4) | R+* | Shall be STATIC. |
| > Radiation Type | (300A,00C6) | R+* | Shall be PHOTON. |
| > High-Dose Technique Type | (300A,00C7) | O+* | If present, must be handled safely |
| > Primary Fluence Mode Sequence | (3002,0050) | R+* |  |
| >> Fluence Mode | (3002,0051) | D |  |
| >> Fluence Mode ID | (3002,0052) | D |  |
| > Treatment Machine Name | (300A,00B2) | R+* | Shall be constant. |
| > Primary Dosimeter Unit | (300A,00B3) | R+ | Shall be MU. |
| > Source-Axis Distance | (300A,00B4) | R+* |  |
| > Beam Limiting Device Sequence | (300A,00B6) | R+* |  |
| >> RT Beam Limiting Device Type | (300A,00B8) | R+* | Shall have at least 2 jaws or at least 1 jaw and 1 MLC. |
| >> Leaf Position Boundaries | (300A,00BE) | R+* | Shall be present for MLCs,. May or may not be present for jaws, may be ignored for jaws |
| > Referenced Patient Setup Number | (300C,006A) | R+* | Shall be >= 1. |
| > Treatment Delivery Type | (300A,00CE) | R+* |  |
| > Number of Wedges | (300A,00D0) | R+* | Shall be 1 or 2. If 2, see also Hard Wedge Beam Modifier |
| > Wedge Sequence | (300A,00D1) | R+* | Shall be present. |
| >> Wedge Type | (300A,00D3) | R+* | Shall be MOTORIZED. Optional Hard Wedge shall be STANDARD |
| >> Wedge ID | (300A,00D4) | R+ |  |
| >> Wedge Angle | (300A,00D5) | RC+ | Shall be present if Wedge Type (300A,00D3) is STANDARD. May be present otherwise. |
| >> Wedge Orientation | (300A,00D8) | R+ |  |
| >> Source to Wedge Tray Distance | (300A,00DA) | RC+ | Shall be present if Wedge Type (300A,00D3) is STANDARD. |
| >>Effective Wedge Angle | (300A,00DE) | -/RC+ RC+/O+* | Shall be present if Wedge Type (300A,00D3) is MOTORIZED. A TMS Actor is required to consume and process this value. A beam producer/consumer actor (e.g., a TPS) may consume this value and is required to produce it |
| > Number of Compensators | (300A,00E0) | R+* | Shall be 0 or 1. If 1, see Compensator Beam Modifier. |
| > Number of Boli | (300A, 00ED) | R+* | Shall be >= 0. If > 0, see Bolus Beam Modifier. |
| > Number of Blocks | (300A,00F0) | R+* | Shall be 0-8. If > 0, see Block Beam Modifier. |
| > Applicator Sequence | (300A,0107) | R+* | Shall not be present. |
| > Final Cumulative Meterset Weight | (300A,010E) | R+* |  |
| > Number of Control Points | (300A,0110) | R+* | Shall be 4. |
| > Control Point Sequence | (300A,0111) | R+* |  |
| >> Cumulative Meterset Weight | (300A,0134) | R+ |  |
| >> Referenced Dose Reference Sequence | (300C,0050) | -/R+* R+/O+* | A TMS Actor is required to consume and process this value. A beam producer/consumer actor (e.g., a TPS) may consume this value and is required to produce it. Shall have at least one item for target dose accumulation. |
| >>> Cumulative Dose Reference Coefficient | (300A,010C) | R+* | Shall be present. |
| >> Nominal Beam Energy | (300A,0114) | R+ | Shall be constant. |
| >> Dose Rate Set | (300A,0115) | R+ | Shall be constant. |
| >> Wedge Position Sequence | (300A,0116) | R+* | Shall be present and consistent with the Wedge Sequence (300A,00D1). |
| >>> Wedge Position | (300A,0118) | R+* | For Motorized Wedge, shall be IN for CPs 0 and 1, OUT for CPs 2 and 3. Shall be IN for optional Hard Wedge |
| >> Beam Limiting Device Position Sequence | (300A,011A) | R+* | Shall be consistent with the Beam Limiting Device Sequence (300A,00B6). |
| >>>Leaf/Jaw Positions | (300A,011C) | R+* |  |
| >> Gantry Angle | (300A,011E) | R+* | Shall be constant. |
| >> Gantry Rotation Direction | (300A,011F) | R+* | Shall be NONE. |
| >> Gantry Pitch Angle | (300A,014A) | O+* | If not present, shall be assumed to be in the zero position. If present, shall be zero. |
| >> Gantry Pitch Rotation Direction | (300A,014C) | O+* | If present, shall be NONE. |
| >> Beam Limiting Device Angle | (300A,0120) | R+* | Shall be constant. |
| >> Beam Limiting Device Rotation Direction | (300A,0121) | R+* | Shall be NONE. |
| < Insert Control Point Sequence Fixed Attributes List > (See 7.4.4.2.1) |  |  |  |
| >> Isocenter Position | (300A,012C) | R+ | Shall be constant for all CPs. |

##### RT Beams Module for Basic Static Electron Beam

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Module Definition

| Attribute | Tag | Beam Technique |  |
| --- | --- | --- | --- |
|  |  | Basic Static Electron |  |
|  |  | Presence | Specific Rules |
| Beam Sequence | (300A,00B0) | R+* |  |
| > Beam Number | (300A,00C0) | R+* | Shall be >= 1. |
| > Beam Name | (300A,00C2) | R+ |  |
| > Beam Type | (300A,00C4) | R+* | Shall be STATIC. |
| > Radiation Type | (300A,00C6) | R+* | Shall be ELECTRON. |
| > High-Dose Technique Type | (300A,00C7) | O+* | If present, must be handled safely |
| > Primary Fluence Mode Sequence | (3002,0050) | R+* |  |
| >> Fluence Mode | (3002,0051) | D |  |
| >> Fluence Mode ID | (3002,0052) | D |  |
| > Treatment Machine Name | (300A,00B2) | R+* | Shall be constant. |
| > Primary Dosimeter Unit | (300A,00B3) | R+ | Shall be MU. |
| > Source-Axis Distance | (300A,00B4) | R+* |  |
| > Beam Limiting Device Sequence | (300A,00B6) | R+* |  |
| >> RT Beam Limiting Device Type | (300A,00B8) | R+* | Shall be 2 jaws, MLC shall not be present |
| >> Leaf Position Boundaries | (300A,00BE) | O+* | NA (no MLC) May or may not be present for jaws, may be ignored for jaws |
| > Referenced Patient Setup Number | (300C,006A) | R+* | Shall be >= 1. |
| > Treatment Delivery Type | (300A,00CE) | R+* |  |
| > Number of Wedges | (300A,00D0) | R+* | Shall be 0. |
| > Number of Compensators | (300A,00E0) | R+* | Shall be 0 or 1. If 1, see Compensator Beam Modifier. |
| > Number of Boli | (300A, 00ED) | R+* | Shall be >= 0. If > 0, see Bolus Beam Modifier. |
| > Number of Blocks | (300A,00F0) | R+* | Shall be 0-8. If > 0, see Block Beam Modifier. |
| > Applicator Sequence | (300A,0107) | R+* | Shall contain 1 item. |
| >> Applicator ID | (300A,0108) | R+ |  |
| >> Applicator Type | (300A,0109) | R+* |  |
| >> Applicator Geometry Sequence | (300A,0431) | R+* |  |
| > Final Cumulative Meterset Weight | (300A,010E) | R+* |  |
| > Number of Control Points | (300A,0110) | R+* | Shall be 2. |
| > Control Point Sequence | (300A,0111) | R+* |  |
| >> Cumulative Meterset Weight | (300A,0134) | R+ |  |
| >> Referenced Dose Reference Sequence | (300C,0050) | -/R+* R+/O+* | A TMS Actor is required to consume and process this value. A beam producer/consumer actor (e.g., a TPS) may consume this value and is required to produce it. Shall have at least one item for target dose accumulation. |
| >>> Cumulative Dose Reference Coefficient | (300A,010C) | R+* | Shall be present. |
| >> Nominal Beam Energy | (300A,0114) | R+ | Shall be constant. |
| >> Dose Rate Set | (300A,0115) | R+ | Shall be constant. |
| >> Wedge Position Sequence | (300A,0116) | R+* | Shall not be present |
| >> Beam Limiting Device Position Sequence | (300A,011A) | R+* | Shall be consistent with the Beam Limiting Device Sequence (300A,00B6). |
| >>>Leaf/Jaw Positions | (300A,011C) | R+* |  |
| >> Gantry Angle | (300A,011E) | R+* | Shall be constant. |
| >> Gantry Rotation Direction | (300A,011F) | R+* | Shall be NONE. |
| >> Gantry Pitch Angle | (300A,014A) | O+* | If not present, shall be assumed to be in the zero position. If present, shall be zero. |
| >> Gantry Pitch Rotation Direction | (300A,014C) | O+* | If present, shall be NONE. |
| >> Beam Limiting Device Angle | (300A,0120) | R+* | Shall be constant. |
| >> Beam Limiting Device Rotation Direction | (300A,0121) | R+* | Shall be NONE. |
| < Insert Control Point Sequence Fixed Attributes List > (See 7.4.4.2.1) |  |  |  |
| >> Isocenter Position | (300A,012C) | R+ | Shall be constant for all CPs. |

##### RT Beams Module for Step & Shoot Beam

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Module Definition

| Attribute | Tag | Beam Technique |  |
| --- | --- | --- | --- |
|  |  | Step & Shoot |  |
|  |  | Presence | Specific Rules |
| Beam Sequence | (300A,00B0) | R+* |  |
| > Beam Number | (300A,00C0) | R+* | Shall be >= 1. |
| > Beam Name | (300A,00C2) | R+ |  |
| > Beam Type | (300A,00C4) | R+* | Shall be STATIC. |
| > Radiation Type | (300A,00C6) | R+* | Shall be PHOTON. |
| > High-Dose Technique Type | (300A,00C7) | O+* | If present, must be handled safely |
| > Primary Fluence Mode Sequence | (3002,0050) | R+* |  |
| >> Fluence Mode | (3002,0051) | D |  |
| >> Fluence Mode ID | (3002,0052) | D |  |
| > Treatment Machine Name | (300A,00B2) | R+* | Shall be constant. |
| > Primary Dosimeter Unit | (300A,00B3) | R+ | Shall be MU. |
| > Source-Axis Distance | (300A,00B4) | R+* |  |
| > Beam Limiting Device Sequence | (300A,00B6) | R+* |  |
| >> RT Beam Limiting Device Type | (300A,00B8) | R+* | At least 1 MLC shall be present |
| >> Leaf Position Boundaries | (300A,00BE) | R+* | Shall be present for MLCs. May or may not be present for jaws, may be ignored for jaws |
| > Referenced Patient Setup Number | (300C,006A) | R+* | Shall be >= 1. |
| > Treatment Delivery Type | (300A,00CE) | R+* |  |
| > Number of Wedges | (300A,00D0) | R+* | Shall be 0 or 1. If 1, see Hard Wedge Beam Modifier |
| > Number of Compensators | (300A,00E0) | R+* | Shall be 0.. |
| > Number of Boli | (300A, 00ED) | R+* | Shall be >= 0. If > 0, see Bolus Beam Modifier. |
| > Number of Blocks | (300A,00F0) | R+* | Shall be 0-8. If > 0, see Block Beam Modifier. |
| > Applicator Sequence | (300A,0107) | R+* | Shall not be present. |
| > Final Cumulative Meterset Weight | (300A,010E) | R+* |  |
| > Number of Control Points | (300A,0110) | R+* | Shall be 2n, where n is the number of unique field shapes composing the beam If the Consumer has a limit, it must document this and safely handle input that exceeds the limit |
| > Control Point Sequence | (300A,0111) | R+* |  |
| >> Cumulative Meterset Weight | (300A,0134) | R+ | CP[0]=0.0 :: CP[2n + 1]=Cumulative Meterset Weight after completion of delivery of the field shape :: CP[2n+1] = CP[2n + 2] |
| >> Referenced Dose Reference Sequence | (300C,0050) | -/R+* R+/O+* | A TMS Actor is required to consume and process this value. A beam producer/consumer actor (e.g., a TPS) may consume this value and is required to produce it. Shall have at least one item for target dose accumulation. |
| >>> Cumulative Dose Reference Coefficient | (300A,010C) | R+* | Shall be present. |
| >> Nominal Beam Energy | (300A,0114) | R+ | Shall be constant. |
| >> Dose Rate Set | (300A,0115) | R+ | Shall be constant. |
| >> Wedge Position Sequence | (300A,0116) | R+* | See Hard Wedge Beam Modifier If present, may not be ignored |
| >>> Wedge Position | (300A,0118) | R+* | Shall be IN. |
| >> Beam Limiting Device Position Sequence | (300A,011A) | R+* | Shall be consistent with the Beam Limiting Device Sequence (300A,00B6). |
| >>>Leaf/Jaw Positions | (300A,011C) | R+* |  |
| >> Gantry Angle | (300A,011E) | R+* | Shall be constant. |
| >> Gantry Rotation Direction | (300A,011F) | R+* | Shall be NONE. |
| >> Gantry Pitch Angle | (300A,014A) | O+* | If not present, shall be assumed to be in the zero position. If present, shall be zero. |
| >> Gantry Pitch Rotation Direction | (300A,014C) | O+* | If present, shall be NONE. |
| >> Beam Limiting Device Angle | (300A,0120) | R+* | Shall be constant. |
| >> Beam Limiting Device Rotation Direction | (300A,0121) | R+* | Shall be NONE. |
| < Insert Control Point Sequence Fixed Attributes List > (See 7.4.4.2.1) |  |  |  |
| >> Isocenter Position | (300A,012C) | R+ | Shall be constant for all CPs. |

##### RT Beams Module for Sliding Window Beam

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Module Definition

| Attribute | Tag | Beam Technique |  |
| --- | --- | --- | --- |
|  |  | Sliding Window |  |
|  |  | Presence | Specific Rules |
| Beam Sequence | (300A,00B0) | R+* |  |
| > Beam Number | (300A,00C0) | R+* | Shall be >= 1. |
| > Beam Name | (300A,00C2) | R+ |  |
| > Beam Type | (300A,00C4) | R+* | Shall be DYNAMIC. |
| > Radiation Type | (300A,00C6) | R+* | Shall be PHOTON. |
| > High-Dose Technique Type | (300A,00C7) | O+* | If present, must be handled safely |
| > Primary Fluence Mode Sequence | (3002,0050) | R+* |  |
| >> Fluence Mode | (3002,0051) | D |  |
| >> Fluence Mode ID | (3002,0052) | D |  |
| > Treatment Machine Name | (300A,00B2) | R+* | Shall be constant. |
| > Primary Dosimeter Unit | (300A,00B3) | R+ | Shall be MU. |
| > Source-Axis Distance | (300A,00B4) | R+* |  |
| > Beam Limiting Device Sequence | (300A,00B6) | R+* |  |
| >> RT Beam Limiting Device Type | (300A,00B8) | R+* | At least 1 MLC shall be present |
| >> Leaf Position Boundaries | (300A,00BE) | R+* | Shall be present for MLCs. May or may not be present for jaws, may be ignored for jaws |
| > Referenced Patient Setup Number | (300C,006A) | R+* | Shall be >= 1. |
| > Treatment Delivery Type | (300A,00CE) | R+* |  |
| > Number of Wedges | (300A,00D0) | R+* | Shall be 0 or 1. If 1, see Hard Wedge Beam Modifier |
| > Number of Compensators | (300A,00E0) | R+* | Shall be 0.. |
| > Number of Boli | (300A, 00ED) | R+* | Shall be >= 0. If > 0, see Bolus Beam Modifier. |
| > Number of Blocks | (300A,00F0) | R+* | Shall be 0-8. If > 0, see Block Beam Modifier. |
| > Applicator Sequence | (300A,0107) | R+* | Shall not be present. |
| > Final Cumulative Meterset Weight | (300A,010E) | R+* |  |
| > Number of Control Points | (300A,0110) | R+* | Shall >2. If the Consumer has a limit, it must document this and safely handle input that exceeds the limit |
| > Control Point Sequence | (300A,0111) | R+* |  |
| >> Cumulative Meterset Weight | (300A,0134) | R+ |  |
| >> Referenced Dose Reference Sequence | (300C,0050) | -/R+* R+/O+* | A TMS Actor is required to consume and process this value. A beam producer/consumer actor (e.g., a TPS) may consume this value and is required to produce it. Shall have at least one item for target dose accumulation. |
| >>> Cumulative Dose Reference Coefficient | (300A,010C) | R+* | Shall be present. |
| >> Nominal Beam Energy | (300A,0114) | R+ | Shall be constant. |
| >> Dose Rate Set | (300A,0115) | R+ | Shall be constant. |
| >> Wedge Position Sequence | (300A,0116) | R+* | See Hard Wedge Beam Modifier If present, may not be ignored |
| >>> Wedge Position | (300A,0118) | R+* | Shall be IN. |
| >> Beam Limiting Device Position Sequence | (300A,011A) | R+* | Shall be consistent with the Beam Limiting Device Sequence (300A,00B6). |
| >>>Leaf/Jaw Positions | (300A,011C) | R+* |  |
| >> Gantry Angle | (300A,011E) | R+* | Shall be constant. |
| >> Gantry Rotation Direction | (300A,011F) | R+* | Shall be NONE. |
| >> Gantry Pitch Angle | (300A,014A) | O+* | If not present, shall be assumed to be in the zero position. If present, shall be zero. |
| >> Gantry Pitch Rotation Direction | (300A,014C) | O+* | If present, shall be NONE. |
| >> Beam Limiting Device Angle | (300A,0120) | R+* | Shall be constant. |
| >> Beam Limiting Device Rotation Direction | (300A,0121) | R+* | Shall be NONE. |
| < Insert Control Point Sequence Fixed Attributes List > (See 7.4.4.2.1) |  |  |  |
| >> Isocenter Position | (300A,012C) | R+ | Shall be constant for all CPs. |

##### RT Beams Module for IMAT/VMAT Beam

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Module Definition

| Attribute | Tag | Beam Technique |  |
| --- | --- | --- | --- |
|  |  | IMAT/VMAT |  |
|  |  | Presence | Specific Rules |
| Beam Sequence | (300A,00B0) | R+* |  |
| > Beam Number | (300A,00C0) | R+* | Shall be >= 1. |
| > Beam Name | (300A,00C2) | R+ |  |
| > Beam Type | (300A,00C4) | R+* | Shall be DYNAMIC. |
| > Radiation Type | (300A,00C6) | R+* | Shall be PHOTON. |
| > High-Dose Technique Type | (300A,00C7) | O+* | If present, must be handled safely |
| > Primary Fluence Mode Sequence | (3002,0050) | R+* |  |
| >> Fluence Mode | (3002,0051) | D |  |
| >> Fluence Mode ID | (3002,0052) | D |  |
| > Treatment Machine Name | (300A,00B2) | R+* | Shall be constant. |
| > Primary Dosimeter Unit | (300A,00B3) | R+ | Shall be MU. |
| > Source-Axis Distance | (300A,00B4) | R+* |  |
| > Beam Limiting Device Sequence | (300A,00B6) | R+* |  |
| >> RT Beam Limiting Device Type | (300A,00B8) | R+* | At least 1 MLC shall be present |
| >> Leaf Position Boundaries | (300A,00BE) | R+* | Shall be present for MLCs. May or may not be present for jaws, may be ignored for jaws |
| > Referenced Patient Setup Number | (300C,006A) | R+* | Shall be >= 1. |
| > Treatment Delivery Type | (300A,00CE) | R+* |  |
| > Number of Wedges | (300A,00D0) | R+* | Shall be 0. |
| > Number of Compensators | (300A,00E0) | R+* | Shall be 0.. |
| > Number of Boli | (300A, 00ED) | R+* | Shall be >= 0. If > 0, see Bolus Beam Modifier. |
| > Number of Blocks | (300A,00F0) | R+* | Shall be 0. |
| > Applicator Sequence | (300A,0107) | R+* | Shall not be present. |
| > Final Cumulative Meterset Weight | (300A,010E) | R+* |  |
| > Number of Control Points | (300A,0110) | R+* | Shall be > 2. If the Consumer has a limit, it must document this and safely handle input that exceeds the limit |
| > Control Point Sequence | (300A,0111) | R+* |  |
| >> Cumulative Meterset Weight | (300A,0134) | R+ |  |
| >> Referenced Dose Reference Sequence | (300C,0050) | -/R+* R+/O+* R+* | A TMS Actor is required to consume and process this value. A beam producer/consumer actor (e.g., a TPS) may consume this value and is required to produce it. Shall have at least one item for target dose accumulation. |
| >>> Cumulative Dose Reference Coefficient | (300A,010C) | R+* | Shall be present. |
| >> Nominal Beam Energy | (300A,0114) | R+ | Shall be constant. |
| >> Dose Rate Set | (300A,0115) | R+ | Shall be present as Nominal Dose Rate. |
| >> Wedge Position Sequence | (300A,0116) | R+* | Shall not be present. |
| >> Beam Limiting Device Position Sequence | (300A,011A) | R+* | Shall be consistent with the Beam Limiting Device Sequence (300A,00B6). |
| >>>Leaf/Jaw Positions | (300A,011C) | R+* |  |
| >> Gantry Angle | (300A,011E) | R+* |  |
| >> Gantry Rotation Direction | (300A,011F) | R+* | Shall be constant (CW or CC) for all CP except the last CP, which can be NONE. |
| >> Gantry Pitch Angle | (300A,014A) | O+* | If not present, shall be assumed to be in the zero position. If present, shall be zero. |
| >> Gantry Pitch Rotation Direction | (300A,014C) | O+* | If present, shall be NONE. |
| >> Beam Limiting Device Angle | (300A,0120) | R+* |  |
| >> Beam Limiting Device Rotation Direction | (300A,0121) | R+* | If present, shall not be ignored. |
| < Insert Control Point Sequence Fixed Attributes List > (See 7.4.4.2.1) |  |  |  |
| >> Isocenter Position | (300A,012C) | R+ | Shall be constant for all CPs. |

##### RT Beams Module for Photon Applicator Beam

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Module Definition

| Attribute | Tag | Beam Technique |  |
| --- | --- | --- | --- |
|  |  | Photon Applicator |  |
|  |  | Presence | Specific Rules |
| Beam Sequence | (300A,00B0) | R+* |  |
| > Beam Number | (300A,00C0) | R+* | Shall be >= 1. |
| > Beam Name | (300A,00C2) | R+ |  |
| > Beam Type | (300A,00C4) | R+* | Shall be STATIC. |
| > Radiation Type | (300A,00C6) | R+* | Shall be PHOTON. |
| > High-Dose Technique Type | (300A,00C7) | O+* | If present, must be handled safely |
| > Primary Fluence Mode Sequence | (3002,0050) | R+* |  |
| >> Fluence Mode | (3002,0051) | D |  |
| >> Fluence Mode ID | (3002,0052) | D |  |
| > Treatment Machine Name | (300A,00B2) | R+* | Shall be constant. |
| > Primary Dosimeter Unit | (300A,00B3) | R+ | Shall be MU. |
| > Source-Axis Distance | (300A,00B4) | R+* |  |
| > Beam Limiting Device Sequence | (300A,00B6) | R+* |  |
| >> RT Beam Limiting Device Type | (300A,00B8) | R+* | Shall be 2 jaws, MLC shall not be present |
| >> Leaf Position Boundaries | (300A,00BE) | O+* | NA (no MLC) May or may not be present for jaws, may be ignored for jaws |
| > Referenced Patient Setup Number | (300C,006A) | R+* | Shall be >= 1. |
| > Treatment Delivery Type | (300A,00CE) | R+* |  |
| > Number of Wedges | (300A,00D0) | R+* | Shall be 0. |
| > Number of Compensators | (300A,00E0) | R+* | Shall be 0. |
| > Number of Boli | (300A, 00ED) | R+* | Shall be >= 0. If > 0, see Bolus Beam Modifier. |
| > Number of Blocks | (300A,00F0) | R+* | Shall be 0. |
| > Applicator Sequence | (300A,0107) | R+* | Shall contain 1 item. |
| >> Applicator ID | (300A,0108) | R+ |  |
| >> Applicator Type | (300A,0109) | R+* | Shall be PHOTON_CIRC |
| >> Applicator Geometry Sequence | (300A,0431) | R+* |  |
| >> Applicator Aperture Shape | (300A,0432) | R+ | Shall be SYM_CIRCULAR |
| > Final Cumulative Meterset Weight | (300A,010E) | R+* |  |
| > Number of Control Points | (300A,0110) | R+* | Shall be 2. |
| > Control Point Sequence | (300A,0111) | R+* |  |
| >> Cumulative Meterset Weight | (300A,0134) | R+ |  |
| >> Referenced Dose Reference Sequence | (300C,0050) | -/R+* R+/O+* | A TMS Actor is required to consume and process this value. A beam producer/consumer actor (e.g., a TPS) may consume this value and is required to produce it. Shall have at least one item for target dose accumulation. |
| >>> Cumulative Dose Reference Coefficient | (300A,010C) | R+* | Shall be present. |
| >> Nominal Beam Energy | (300A,0114) | R+ | Shall be constant. |
| >> Dose Rate Set | (300A,0115) | R+ | Shall be constant. |
| >> Wedge Position Sequence | (300A,0116) | R+* | Shall not be present |
| >> Beam Limiting Device Position Sequence | (300A,011A) | R+* | Shall be consistent with the Beam Limiting Device Sequence (300A,00B6). |
| >>>Leaf/Jaw Positions | (300A,011C) | R+* |  |
| >> Gantry Angle | (300A,011E) | R+* | Shall be constant. |
| >> Gantry Rotation Direction | (300A,011F) | R+* | Shall be NONE. |
| >> Gantry Pitch Angle | (300A,014A) | O+* | If not present, shall be assumed to be in the zero position. If present, shall be zero. |
| >> Gantry Pitch Rotation Direction | (300A,014C) | O+* | If present, shall be NONE. |
| >> Beam Limiting Device Angle | (300A,0120) | R+* | Shall be constant. |
| >> Beam Limiting Device Rotation Direction | (300A,0121) | R+* | Shall be NONE. |
| < Insert Control Point Sequence Fixed Attributes List > (See 7.4.4.2.1) |  |  |  |
| >> Isocenter Position | (300A,012C) | R+ | Shall be constant for all CPs. |

##### RT Beams Module for Photon Applicator Arc Beam

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Module Definition

| Attribute | Tag | Beam Technique |  |
| --- | --- | --- | --- |
|  |  | Photon Applicator Arc |  |
|  |  | Presence | Specific Rules |
| Beam Sequence | (300A,00B0) | R+* |  |
| > Beam Number | (300A,00C0) | R+* | Shall be >= 1. |
| > Beam Name | (300A,00C2) | R+ |  |
| > Beam Type | (300A,00C4) | R+* | Shall be DYNAMIC. |
| > Radiation Type | (300A,00C6) | R+* | Shall be PHOTON. |
| > High-Dose Technique Type | (300A,00C7) | O+* | If present, must be handled safely |
| > Primary Fluence Mode Sequence | (3002,0050) | R+* |  |
| >> Fluence Mode | (3002,0051) | D |  |
| >> Fluence Mode ID | (3002,0052) | D |  |
| > Treatment Machine Name | (300A,00B2) | R+* | Shall be constant. |
| > Primary Dosimeter Unit | (300A,00B3) | R+ | Shall be MU. |
| > Source-Axis Distance | (300A,00B4) | R+* |  |
| > Beam Limiting Device Sequence | (300A,00B6) | R+* |  |
| >> RT Beam Limiting Device Type | (300A,00B8) | R+* | Shall be 2 jaws, MLC shall not be present |
| >> Leaf Position Boundaries | (300A,00BE) | O+* | NA (no MLC) May or may not be present for jaws, may be ignored for jaws |
| > Referenced Patient Setup Number | (300C,006A) | R+* | Shall be >= 1. |
| > Treatment Delivery Type | (300A,00CE) | R+* |  |
| > Number of Wedges | (300A,00D0) | R+* | Shall be 0. |
| > Number of Compensators | (300A,00E0) | R+* | Shall be 0. |
| > Number of Boli | (300A, 00ED) | R+* | Shall be >= 0. If > 0, see Bolus Beam Modifier. |
| > Number of Blocks | (300A,00F0) | R+* | Shall be 0. |
| > Applicator Sequence | (300A,0107) | R+* | Shall contain 1 item. |
| >> Applicator ID | (300A,0108) | R+ |  |
| >> Applicator Type | (300A,0109) | R+* | Shall be PHOTON_CIRC. |
| >> Applicator Geometry Sequence | (300A,0431) | R+* |  |
| >> Applicator Aperture Shape | (300A,0432) | R+ | Shall be SYM_CIRCULAR |
| > Final Cumulative Meterset Weight | (300A,010E) | R+* |  |
| > Number of Control Points | (300A,0110) | R+* | Shall be 2. |
| > Control Point Sequence | (300A,0111) | R+* |  |
| >> Cumulative Meterset Weight | (300A,0134) | R+ |  |
| >> Referenced Dose Reference Sequence | (300C,0050) | -/R+* R+/O+* | A TMS Actor is required to consume and process this value. A beam producer/consumer actor (e.g., a TPS) may consume this value and is required to produce it. Shall have at least one item for target dose accumulation. |
| >>> Cumulative Dose Reference Coefficient | (300A,010C) | R+* | Shall be present. |
| >> Nominal Beam Energy | (300A,0114) | R+ | Shall be constant. |
| >> Dose Rate Set | (300A,0115) | R+ | Shall be constant. |
| >> Wedge Position Sequence | (300A,0116) | R+* | Shall not be present |
| >> Beam Limiting Device Position Sequence | (300A,011A) | R+* | Shall be consistent with the Beam Limiting Device Sequence (300A,00B6). |
| >>>Leaf/Jaw Positions | (300A,011C) | R+* |  |
| >> Gantry Angle | (300A,011E) | R+* |  |
| >> Gantry Rotation Direction | (300A,011F) | R+* | Shall be CW or CC for Control Point 0 Can be NONE for Control Point 1. |
| >> Gantry Pitch Angle | (300A,014A) | O+* | If not present, shall be assumed to be in the zero position. If present, shall be zero. |
| >> Gantry Pitch Rotation Direction | (300A,014C) | O+* | If present, shall be NONE. |
| >> Beam Limiting Device Angle | (300A,0120) | R+* | Shall be constant. |
| >> Beam Limiting Device Rotation Direction | (300A,0121) | R+* | Shall be NONE. |
| < Insert Control Point Sequence Fixed Attributes List > (See 7.4.4.2.1) |  |  |  |
| >> Isocenter Position | (300A,012C) | R+ | Shall be constant for all CPs. |

#### General Beam Attribute Specifications

##### Control Point Fixed Attribute List Base Content

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Required Attributes

The list of attributes and requirements below shall be included in all TPPC transactions as noted in the RT Beam Module specification associated with those transactions.

| Attribute |  | Tag | Control Point Sequence |  |
| --- | --- | --- | --- | --- |
|  |  |  | Fixed Attributes |  |
|  |  |  | Presence | Specific Rules |
| >> Patient Support Angle | (300A,0122) | R+* | Shall be constant. |
| >> Patient Support Rotation Direction | (300A,0123) | R+* | Shall be NONE. |
| >> Table Top Eccentric Axis Distance | (300A,0124) | O+* | If present, shall be constant. |
| >> Table Top Eccentric Angle | (300A,0125) | R+* | Shall be zero. |
| >>Table Top Eccentric Rotation Direction | (300A,0126) | R+* | Shall be NONE. |
| >> Table Top Pitch Angle | (300A,0140) | R+* | Shall be zero. |
| >> Table Top Pitch Rotation Direction | (300A,0142) | R+* | Shall be NONE. |
| >> Table Top Roll Angle | (300A,0144) | R+* | Shall be zero. |
| >> Table Top Roll Rotation Direction | (300A,0146) | R+* | Shall be NONE |
| >> Table Top Vertical Position | (300A,0128) | O+* | If value is present, shall be constant. |
| >> Table Top Longitudinal Position | (300A,0129) | O+* | If value is present, shall be constant. |
| >> Table Top Lateral Position | (300A,012A) | O+* | If value is present, shall be constant. |

#### Beam Option Specifications

##### Bolus Beam Modifier Base Content

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Required Attributes

One or more Boli may be optionally included in any of the Treatment Planning - Plan Content Storage and Retrieval transactions (Producer and Consumer actors).

For actors supporting the Bolus Beam Modifier, the attributes specified in the table below have these additional requirements if Number of Boli (300A,00ED) is greater than zero.

| Attribute | Tag | Beam Modifier |  |
| --- | --- | --- | --- |
|  |  | Bolus |  |
|  |  | Presence | Specific Rules |
| > Number of Boli | (300A,00ED) | R+* | Shall be >=1. |
| > Referenced Bolus Sequence | (300A,00B0) | R+* |  |
| >> Bolus ID | (300A,00DC) | R+* | Shall be present. |

##### Block Beam Modifier Base Content

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Required Attributes

One or more Blocks may be optionally included in some of the Treatment Planning - Plan Content Storage and Retrieval transactions (Producer and Consumer actors):

For actors supporting the Block Beam Modifier, the attributes specified in the table below have these additional requirements if Number of Blocks (300A,00F0) is greater than zero.

| Attribute | Tag | Beam Modifier |  |
| --- | --- | --- | --- |
|  |  | Block |  |
|  |  | Presence | Specific Rules |
| > Number of Blocks | (300A,00F0) | R+* | Photon Beams: Shall be 0 - 8. Electron Beams: Shall be 0 or 1. |
| > Block Sequence | (300A,00F4) | R+* |  |
| >> Block Tray ID | (300A,00F5) | R+ | See Note 1. |
| >> Source to Block Tray Distance | (300A,00F6) | R+ |  |
| >> Block Divergence | (300A,00FA) | R+* |  |
| >> Block Mounting Position | (300A,00FB) | R+ | Shall be present, and shall be handled safely for enumerated values not supported. |
| >> Material ID | (300A,00E1) | R+ |  |
| >> Block Thickness | (300A,0100) | R+ |  |
| >> Block Number of Points | (300A,0104) | R+* |  |
| >> Block Data | (300A,0106) | R+* |  |

Note 1:

Photon Beams: There may be multiple blocks with the same Block Tray ID (i.e., placed on the same Tray, e.g., an aperture block and a shield block). If/when Treatment Delivery Verification is taking place, it is essentially the Block Tray that is verified, not the individual blocks on the tray, so one would expect the same Accessory ID to be reused for the same Block Tray ID. This is why Block Name is not made Mandatory (R+*) for the profile. Only a single Block Tray ID shall be supported for a given beam.

Electron Beams: The Block Tray ID defines the electron insert which is checked by the TDD.

##### Compensator Beam Modifier Base Content

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Required Attributes

A single Compensator may be optionally included in some of the Treatment Planning - Plan Content Storage and Retrieve transactions (Producer and Consumer actors):

For actors supporting the Compensator Beam Modifier, the attributes specified in the table below have these additional requirements: if Number of Compensators (300A,00E0) is equal to one.

| Attribute | Tag | Beam Modifier |  |
| --- | --- | --- | --- |
|  |  | Compensator |  |
|  |  | Presence | Specific Rules |
| Beam Sequence | (300A,00B0) | R+* |  |
| > Number of Compensators | (300A,00E0) | R+* | Shall be 1. |
| > Compensator Sequence | (300A,00E3) | R+* |  |
| >> Compensator Type | (300A,00EE) | R+* | Shall be STANDARD. |
| >> Material ID | (300A,00E1) | R+* |  |
| >> Compensator ID | (300A,00E5) | R+* |  |
| >> Source to Compensator Tray Distance | (300A,00E6) | R+* |  |
| >> Compensator Divergence | (300A,02E0) | R+* |  |
| >> Compensator Mounting Position | (300A,02E1) | R+* | Shall be PATIENT_SIDE or SOURCE_SIDE. |
| >> Compensator Transmission Data | (300A,00EB) | R+* |  |
| >> Compensator Thickness Data | (300A,00EC) | R+* |  |

##### Hard Wedge Beam Modifier Base Content

###### Referenced Standards

DICOM 2015a Edition PS 3.3

###### Required Attributes

A single Hard Wedge may be optionally included in some of the Treatment Planning - Plan Content Storage and Retrieve transactions (Producer and Consumer actors):

For actors supporting the Hard Wedge Beam Modifier, the attributes specified in the table below have these additional requirements if  Number of Wedges (300A,00D0) is greater than 0.

| Attribute | Tag | Beam Modifier |  |
| --- | --- | --- | --- |
|  |  | Hard Wedge |  |
|  |  | Presence | Specific Rules |
| Beam Sequence | (300A,00B0) | R+* |  |
| > Number of Wedges | (300A,00D0) | R+* | Shall be 1 or 2; if 2, one will be Hard and the other will be either Motorized or Virtual |
| > Wedge Sequence | (300A,00D1) | R+* |  |
| >> Wedge Type | (300A,00D3) | R+* | Shall be STANDARD. |
| >> Wedge ID | (300A,00D4) | R+ |  |
| >> Wedge Angle | (300A,00D5) | R+ |  |
| >> Wedge Orientation | (300A,00D8) | R+ |  |
| >> Source to Wedge Tray Distance | (300A,00DA) | R+ |  |
| … |  |  |  |
| > Control Point Sequence | (300A,0111) | R+* |  |
| >> Wedge Position Sequence | (300A,0116) | R+* | If present may not be ignored. |
| >>> Wedge Position | (300A,0118) | R+* | Shall be IN. |