IHE Technical Framework Supplement - <Profile Name (Profile Acronym)>
______________________________________________________________________________


Integrating the Healthcare Enterprise

[image: image1.png]



IHE Radiation Oncology
Technical Framework Supplement

Discrete Positioning And Delivery Profile
(DPDW) 

Draft in preparation for Public Comment

Date:

2011-xx-xx (Version 1.7)
Author:
Ulrich Busch

Email: 

ihe.-ro@ihe.net

Foreword
This is a supplement to the IHE Radiation Oncology Technical Framework 1.0. Each supplement undergoes a process of public comment and trial implementation before being incorporated into the volumes of the Technical Frameworks.
[For Public Comment: This supplement is submitted for Public Comment between  <Month XX, 201X and Month XX, 201X >. Comments are invited and may be submitted to the IHE forums at < http://forums.rsna.org publisher will fill in full url>. In order to be considered in development of the Trial Implementation version of the supplement comments must be received by <Month XX, 201X>. Please use the Public Comment Template provided in the thread and submit comments by posting the completed template as an attachment to a Reply or New Thread.]
[For Trial Implementation: This supplement is submitted for Trial Implementation as of <Month XX, 201X> and will be available for testing at subsequent IHE Connectathons. The supplement may be amended based on the results of testing. Following successful testing it will be incorporated into the <Domain Name> Technical Framework. Comments are invited and may be submitted on the IHE forums at< http://forums.rsna.org publisher will fill in full url> or by email to <domain_name@ihe.net>.]
This supplement describes changes to the existing technical framework documents and where indicated amends text by addition (bold underline) or removal (bold strikethrough), as well as addition of large new sections introduced by editor’s instructions to “add new text” or similar, which for readability are not bolded or underlined.

“Boxed” instructions like the sample below indicate to the Volume Editor how to integrate the relevant section(s) into the relevant Technical Framework volume: 
Replace Section X.X by the following:

General information about IHE can be found at: www.ihe.net
Information about the IHE Radiation Oncology may be found at:  http://www.ihe.net/Domains/index.cfm
Information about the structure of IHE Technical Frameworks and Supplements can be found at: http://www.ihe.net/About/process.cfm and http://www.ihe.net/profiles/index.cfm
The current version of the IHE Technical Framework can be found at: http://www.ihe.net/Technical_Framework/index.cfm
CONTENTS

6Introduction


61.1
Profile Abstract


61.2
History


71.3
Open Issues and Questions


81.4
Closed Issues


91.5
Editorial


10Volume 1 – Integration Profiles


101.6
1.7 History of Annual Changes


101.7
1.n Copyright Permission


101.8
2.1 Dependencies among Integration Profiles


101.8.1
2.2.X Discrete Positioning and Delivery Workflow Integration Profile


112
Treatment Session Workflow Frame Integration Profile


112.1
Actors/ Transactions


122.2
Treatment Session Workflow Frame Integration Profile Options


122.3
Treatment Session Workflow Frame Process Flow


122.3.1
Opening Treatment Session


132.3.2
Closing Treatment Session


152.4
Treatment Session Workflow Frame Security Considerations


163
Discrete Positioning Workflow Integration Profile


163.1
Actors/ Transactions


183.2
Discrete Positioning Workflow Integration Profile Options


183.3
Discrete Positioning Workflow Process Flow


223.4
Discrete Positioning Workflow Security Considerations


234
Discrete Delivery Workflow Integration Profile


234.1
Actors/ Transactions


244.2
Discrete Delivery Workflow Integration Profile Options


254.3
Discrete Delivery Workflow Process Flow


274.4
Discrete Delivery Workflow Security Considerations


285
Discrete Delivery and Monitoring Integration Profile


295.1
Actors/ Transactions


305.2
Discrete Delivery and Monitoring Integration Profile Options


315.3
Discrete Delivery and Monitoring Process Flow


315.3.1
Overall Process Flow


355.3.2
Interruption Handling by Re-Acquisition of Positioning Data


365.3.3
Interruption Handling by Automatic Re-positioning


395.3.4
Interruption Handling by Manual Re-positioning


415.3.5
Interruption Handling by Beam Parameter Adaptation


435.3.6
Interruption Handling by Treatment Termination


435.4
Discrete Delivery and Monitoring Security Considerations


43<Appendix A> Actor Summary Definitions


44<Appendix B> Transaction Summary Definitions


47Glossary


48Volume 2 - Transactions


485.5
RO-DPD-01: Worklist Query for Positioning Acquisition


485.5.1
Scope


485.5.2
Use Case Roles


495.5.3
Referenced Standards


495.5.4
Interaction Diagram


505.6
RO-DPD-02: Retrieve Patient Positioning Device Actuals


505.6.1
Scope


505.6.2
Use Case Roles


515.6.3
Referenced Standards


515.6.4
Interaction Diagram


525.7
RO-DPD-03:  Request RT Patient Position Correction


525.7.1
Scope


525.7.2
Use Case Roles


525.7.3
Referenced standards


535.7.4
Interaction Diagram


545.8
RO-DPD-04:  Store RT Patient Position Correction Instruction


545.8.1
Scope


545.8.2
Use Case Roles


555.8.3
Referenced standards


555.8.4
Interaction Diagram


565.9
RO-DPD-05: Store RT Repositioning Results to Archive


565.9.1
Scope


565.9.2
Use Case Roles


575.9.3
Referenced standards


575.9.4
Interaction Diagram


585.10
RO-DPD-06: Worklist Query for Repositioning


585.10.1
Scope


585.10.2
Use Case Roles


585.10.3
Referenced standards


595.10.4
Interaction Diagram


605.11
RO-DPD-07: Notify on Beam Delivery Status Change


605.11.1
Scope


605.11.2
Use Case Roles


605.11.3
Referenced standards


615.11.4
Interaction Diagram


615.12
RO-DPD-08: Retrieve RT Patient Position Correction Instruction


615.12.1
Scope


625.12.2
Use Case Roles


625.12.3
Referenced standards


635.12.4
Interaction Diagram


645.13
RO-DPD-09: Subscribe / Unsubscribe to Treat UPS


645.13.1
Scope


645.13.2
Use Case Roles


655.13.3
Referenced standards


655.13.4
Interaction Diagram


665.14
RO-DPD-10: Notify on active beam


665.14.1
Scope


665.14.2
Use Case Roles


675.14.3
Referenced standards


675.14.4
Interaction Diagram


685.15
RO-DPD-11: Retrieve Positioning Acquisition Results


685.15.1
Scope


685.15.2
Use Case Roles


685.15.3
Referenced standards


685.15.4
Interaction Diagram


695.16
RO-DPD-12: Worklist Query for Positioning Registration


695.16.1
Scope


695.16.2
Use Case Roles


705.16.3
Referenced Standards


715.16.4
Interaction Diagram


725.17
RO-DPD-13: Worklist Query for Position Monitoring


725.17.1
Scope


725.17.2
Use Case Roles


735.17.3
Referenced Standards


735.17.4
Interaction Diagram


745.18
RO-DPD-14: Retrieve RT Patient Position Modification Instruction


745.18.1
Scope


745.18.2
Use Case Roles


755.18.3
Referenced Standards


755.18.4
Interaction Diagram


765.19
RO-DPD-15: Store RT Patient Position Modification Result


765.19.1
Scope


765.19.2
Use Case Roles


765.19.3
Referenced Standards


775.19.4
Interaction Diagram


775.20
RO-DPD-16: Store Monitoring Results to Archive


775.20.1
Scope


775.20.2
Use Case Roles


785.20.3
Referenced Standards


785.20.4
Interaction Diagram


795.21
RO-DPD-17: Collective UPS Final Update


795.21.1
Scope


795.21.2
Use Case Roles


805.21.3
Referenced Standards


805.21.4
Interaction Diagram


815.22
RO-DPD-18: Collective UPS Completed / Cancelled


815.22.1
Scope


815.22.2
Use Case Roles


825.22.3
Referenced Standards


825.22.4
Interaction Diagram


835.23
RO-DPD-19: Indicate Ready for Monitoring


835.23.1
Scope


835.23.2
Use Case Roles


835.23.3
Referenced Standards


835.23.4
Interaction Diagram


845.24
RO-DPD-20: Notify Device to start UPS


845.24.1
Scope


845.24.2
Use Case Roles


855.24.3
Referenced Standards


855.24.4
Interaction Diagram


885.25
RO-DPD-21: Create Positioning Acquisition and Positioning Registration UPS


885.25.1
Scope


885.25.2
Use Case Roles


885.25.3
Referenced Standard


885.25.4
Interaction Diagram


885.26
RO-DPD-22: Create Treat UPS and Beam Delivery Instruction for Continuation


885.26.1
Scope


885.26.2
Use Case Roles


885.26.3
Referenced Standard


885.26.4
Interaction Diagram


89Appendix General Requirements on using UPS


895.27
General Transaction description


895.28
Matching Keys and Return Keys for Display


915.29
Expected Actions


915.30
Receive Scheduled UWPS Worklist Message


925.31
Trigger Events


925.32
Message Semantics


92<Appendix letter> Appendix_Name




Introduction

This supplement adds the Discrete Positioning and Delivery Workflow Profile to the IHE-RO Domain. 

1.1 Profile Abstract

The Discrete Positioning and Delivery Workflow Profile describes a patient’s treatment in radiation therapy, which involves the positioning of a patient and treatment delivery with or without position monitoring. In this profile, the positioning, monitoring and radiation delivery are executed by a two or more separate devices.

1.2 History

	May 20, 2011
	1.8
	Ulrich Busch
(ulrich.busch@varian.com)
	Various editorial cleanups on formatting, styles etc. (no change bars)

	May 13, 2011
	1.7
	Ulrich Busch
(ulrich.busch@varian.com)
	Added RO-19 to the Initiate Session sequence diagram in the same sense as in IPDW. Added explanatory text as present in IPDW.
RO-DPD-07: Renamed to ‘Notify on Beam Delivery Status Change’ and changed description accordingly to indicate that this is a notification only (and different from RO-26)

	April 15, 2011
	
	Ulrich Busch
(ulrich.busch@varian.com)
	- Split sequence diagram for Acquisition and Registration into ‘Positioning Acquisition’ and ‘Evaluate Position’, because they need more space.

	March 4, 2011
	
	Ulrich Busch
(ulrich.busch@varian.com)
	- In ‘Discrete Delivery Workflow Integration Profile’:

- added transaction RO-DPD-02 from TSM to TDD to prompt TDD to start treating.

- Correction: UPS Final Update [RO-21] and UPS Completed/Cancelled [RO-25] transactions should occur after all beams are treated.

- added RO-DPD-07 right away after WL Query

- In Delivery diagrams, the first RO-DPD-07 is replaced with the a Set UPS IN PROGRESS transaction (like Discrete Delivery Workflow Profile),

- In “Interruption Handling by a Re-Acquisition of Positioning Data,” added treatment record storage.

- Introduced RO-DPD-21: Create Positioning Acquisition and Positioning Registration UPS

- Introduced RO-DPD-22: Create Treat UPS and Beam Delivery Instruction for continuation

- Removed sections ’ 1.1.1 Treatment Completion’ and ‘1.1.2 Profile Preconditions’ which were present in the Open Issues sections to indicate that they need to be moved from the originally and misplaced location. Those sections are now part of the IPDW and therefore don’t need to be placed in here any more – they are referred to from here.

	Jan 27, 2011
	
	Ulrich Busch
(ulrich.busch@varian.com)
	- The RO-DPD-02 Transaction goes from PPVS to PPD (in addition to from PPVS to TDD) in the presence of PPD devices.
- Updated Use Case Roles and Interaction diagrams or RO-DPD-20

- Started to specify in detail how “Notify Device to Start UPS” (RO-DPD-20) is supposed to work. 
- Added safety transaction to scenarios diagram as a option to do (not yet decided, whether we do this or not). 

	Nov 12, 2010
	1.6
	Ulrich Busch
(ulrich.busch@varian.com)
	Converted to Supplement

	May 7, 2010
	1.5
	Ulrich Busch
(ulrich.busch@varian.com)
	New transactions:

Notify Device to start UPS [RO-DPD-20]

Used to notify the esp. devices like PPVS, PPRS, PPMS, PPD that they can start their process

Protocol undefined currently.

Added this transaction at various places, where user initiation was for seen in previous versions.

TDD in re-positioning replaced by PPD Primary

In text

In scenarios

Linked PPD primary and PPD secondary with intiation transaction so that they can synchronize the re-positioning

Assumption is, that due to the complexity of different mechanics and geometric conditions and the required knowledge of all of this, when PPDs work together, that the PPDs regulate the re-positioning amongst each other as needed.

Made RO-14 optional in case a primary PPD can perform all axis which is specified for the a certain delivery device..

Re-arranged some sequences in re-positioning to allow for repetition of re—positioning steps, when both PPD aahve dependent acis.

Added RO-26 progress update from TDD to TMS to ensure treatment progress awareness as safe as possible at the TMS, so that network and application failures are less critical.

Removed Actor combination section and replaced by note, that this shall be re-visited later.

Moved RO-21, RO-26 of PPMS to be performed after all beams are treated.

Moved subscription of PPMS  to TSM, added subscription of TDD

Moved UPS status change transaction of PPMS to TSM, removed UPS Treat Status updates from TDD to PPMS, becaused they go via TSM in this version.

Correction continuation of subsequent beams (and merged with resumption of beam after monitoring interruption)


1.3 Open Issues and Questions

	March 4, 2011
	1.6
	Ulrich Busch
(ulrich.busch@varian.com)
	Support for Positioning Verification and correction on a beam-by-beam basis may have to be included..

	Jan 27, 2011
	1.6
	Ulrich Busch
(ulrich.busch@varian.com)
	Issues of Meeting 2011-01-27:

- Eventually, design the scenario with the safety checker, which works with similar component approach as the other ones – also to illustrate to the approach in a Pull mode.

	June 2010
	Pre- 1.6
	Ulrich Busch
(ulrich.busch@varian.com)
	(Excerpt of remaining items of protocol of meeting 2010-06 :)

DPDW draft document was reviewed by Uli Busch. Changes to to be made in the document are noted below:

•
425 a. TC should evaluate Discrete Positioning Workflow support for positioning-only scenarios in which initial patient setup is performed in a separate room from treatment. In such cases, there may be a separate session manager, and may require transfer of information to a second session manager for treatment.

•
c. Consider including “TDD/Primary PPD” combination Actor

•
e. In section 6.2.5, for dynamic, adaptive therapy, one needs to create a new plan and store a treatment record (alternately report actual in treatment record)


1.4 Closed Issues

	March 4, 2011
	1.6
	Ulrich Busch
(ulrich.busch@varian.com)
	- RO-DPD-07: Change Treat UPS Status : Take similar semantics (and partly same text) as in IPDW: RO-26: UPS Progress Update for Treatment. 

Positioning Verification per beam.

	June 2010
	Pre- 1.6
	Ulrich Busch
(ulrich.busch@varian.com)
	The following iotems have been addressed in version 1.7:

(Excerpt of protocol of meeting 2010-06 :)

DPDW draft document was reviewed by Uli Busch. Changes to to be made in the document are noted below:

•
435 Workflow Integration Profile 1 diagram. a. In figure 5.2-1, add transaction RO-DPD-02 from TSM to TDD to prompt TDD to start treating.

•
b. Correction: UPS Final Update [RO-21] and UPS 440 Completed/Cancelled [RO-25] transactions should occur after all beams are treated.

•
d. Note: UPS in Progress is missing

•
445 e. Note: UPS Progress Update should be synchronized with TDW and IPDW profiles

•
a. Note: Subscription for monitoring treat and other UPSs needs to be clarified in detail with respect to the mechanisms used.

•
450 b. In Fig 6.2-2 Profile 2 diagram, the first RO-DPD-07 should be replaces with the a Set UPS IN PROGRESS transaction (like Discrete Delivery Workflow Profile),

•
c. Correction: RO-DPD-07 should only indicate N-EVENT Report (Beam-on/Beam-off)

•
c. In Section. 6.2-2 “Interruption Handling by a Re-Acquisition of Positioning Data,” add treatment record storage.

•
455 d. At the end of the re-positioning the TSM must issue a continuation UPS and Beam Delivery Instruction for the TDD to continue the treatment.


1.5 Editorial

Diagrams shall be pasted as Picture (Windows Metafile), because that takes the least space.
Volume 1 – Integration Profiles

1.6 1.7 History of Annual Changes

This supplement adds the Discrete Positioning and Delivery Workflow Profile to the IHE-RO Domain. 

Add the following bullet to the end of the bullet list in section 1.7

· Added the Discrete Positioning and Delivery Workflow Profile, which describes a patient’s treatment in radiation therapy, which involves the positioning of a patient and treatment delivery with or without position monitoring. In this profile, the positioning, monitoring and radiation delivery are executed by a two or more separate devices.

1.7 1.n Copyright Permission

<Add information on any standards referenced in the profile that are not already addressed in the permission section.>

Add the following to sections 1.n:

1.8 2.1 Dependencies among Integration Profiles

Add the following to Table 2-1

	<Profile Name>
	<?>
	
<?>
	<->


Add the following section to section 2.2

1.8.1 2.2.X Discrete Positioning and Delivery Workflow Integration Profile

<One paragraph description/ overview of Integration Profile>

Add Section X
2 Treatment Session Workflow Frame Integration Profile

2.1 Actors/ Transactions

<Actor/ Transaction diagram and Actor/ Transaction table. >

Figure X.1-1 shows the actors directly involved in the <Profile Name> Integration Profile and the relevant transactions between them.  Other actors that may be indirectly involved due to their participation in <other related profiles>, etc. are not necessarily shown.


[image: image2.wmf] 

­

 Transaction_3 [3]

 

­

 Transaction_4 [4]

 

 

Actor C

 

Transaction_1 [1] 

¯

 

¯

 Transaction _2 [2]

 

Actor A

 

Actor F

 

Actor D

 

Actor E

 

Actor B

 

Transaction_1 [1] 

­

 

­

 Transaction _2 [2]

 


Figure X.1-1.  <Profile Name> Actor Diagram

Table X.1-1 lists the transactions for each actor directly involved in the <Profile Name> Profile. In order to claim support of this Integration Profile, an implementation must perform the required transactions (labeled “R”). Transactions labeled “O” are optional.  A complete list of options defined by this Integration Profile and that implementations may choose to support is listed in Volume I, Section X.2.

Table X.1-1.  <Profile Name> Integration Profile - Actors and Transactions

	Actors
	Transactions 
	Optionality
	Section in Vol. 2

	TMS
	RO-17 Worklist Query for Positioning and Delivery 
	R
	

	
	RO-19 UPS in Progress
	R
	

	
	RO-20 Retrieve Workitem Input Objects from TMS
	R
	

	TSM
	RO-17 Worklist Query for Positioning and Delivery 
	R
	

	
	RO-19 UPS in Progress
	R
	

	
	RO-20 Retrieve Workitem Input Objects from TMS
	R
	


2.2 Treatment Session Workflow Frame Integration Profile Options

The profile does not contain any options.
2.3 Treatment Session Workflow Frame Process Flow

The initial query, performed by the TSM against the TMS, to retrieve the UPS containing the steps to be performed in the treatment session, has the same content as in the Integrated Positioning and Delivery Workflow Profile (IPDW). In the current profile, the TSM distributes the UPS to the various devices during the treatment session, while in the Integrated Positioning and Delivery Workflow Profile (IPDW) this is an internal communication. In both cases though, the content of the treatment session has been scheduled in a TMS and is retrieved from this schedule to the TSM.

Therefore all general conditions and requirements of the IPDW in the scope of initial and final states of the TMS applies to this profile as well, esp. the content of the sections ‘Treatment Cancellation Prior to Radiation’, ‘Treatment Completion’,  ‘Profile Preconditions’, ‘Profile and Device Capabilities’.
2.3.1 Opening Treatment Session

In all profiles the TSM takes the role of providing a worklist service provider role to the involved devices. The TSM will initially query the TMS to retrieve all UPS planned for this treatment session, when the user indicates that the next patient is to be treated.

This Worklist will contain the Treat UPS and additionally all UPS, which are present to define the positioning and monitoring behavior. Those UPS are published to the actors performing positioning, registration and monitoring on query. The definition of the various possible UPS follows the same rules as for the Integrated Positioning and Delivery Profile (especially along RO-17). The difference of this actual profile to the Integrated Positioning and Delivery Profile is that once the TSM has retrieved all UPS for the treatment session, it distributes those UPS further to other devices as needed in a multi-vendor environment.

After retrieving the UPS, the TSM uses the UPS N-ACTION service to inform the TMS that the specified Unified Procedure Steps have been claimed for execution during the treatment session. This means especially, that the device has taken exclusive ownership of those steps. The start of the actual execution will typically occur later in the course of processing the sequence of the Unified Procedure Steps.

[image: image3.wmf]TMS

Version 

1

.

7

2011

-

05

-

13

TSM

Worklist Query for Positioning and Delivery 

[

RO

-

17

]

Retrieve Workitem Input Object from TMS 

[

RO

-

20

]

User initiates query for Patients to be treated

UPS in Progress  

[

RO

-

19

]


2.3.2 Closing Treatment Session

The various actors working on those UPS update the UPS status by setting the UPS content information and status on the TSM. At latest at the end of the treatment session, the TSM will update all UPS it has on the TMS to communicate back the final result.

[image: image4.wmf]TMS

Version 

1

.

7

2011

-

05

-

13

TSM

Collective UPS Final Update 

[

RO

-

DPD

-

17

]

Collective UPS Completed 

/ 

Cancelled 

[

RO

-

DPD

-

18

]


This outer frame of the interaction follows the sequence diagram as defined in Supplement 74 'Utilization of Worklist in Radiotherapy Treatment Delivery'.


[image: image5.wmf] 

Transaction 2 [2]

 

Transaction 1 [1]

 

Actor D/

 

Actor E

 

Actor A /

 

Actor B

 

Internal 

action1 

 

Transaction

-

D

 

[D]

 

Transaction

-

C [C]

 

Actor B/ Actor C 

 

Transaction 2 [2]

 

Transaction 3 [3]

 

Internal action 

2 

 

Transaction

-

A [A]

 

Transaction

-

B [B]

 

Figure X.2-1. Basic Process Flow in <Profile Name> Profile

2.4 Treatment Session Workflow Frame Security Considerations

TBD

3 Discrete Positioning Workflow Integration Profile

The Discrete Positioning Workflow Profile involves the positioning of a patient prior to treatment delivery, where the positioning and delivery are performed by separate devices subsequently e.g. as follows:

· The Patient Positioning Verification System (PPVS) acquires a for example set of 2D (planar projection), 3D (CT) positioning images or optical markers. This set represents the input data for the following registration against reference data.

· The Patient Positioning Registration System (PPRS) performs a registration of the acquired data against previously retrieved reference images, reference markers positions or similar appropriate reference information. The user than assesses the current patient position with the results of the PPRS and determine the need and the amount of a correction of the patient position.

· Based on the user’s decision, the PPRS requests a repositioning of the patient if necessary. 

· The Patient Positioning Devices (PPD) performs the re-positioning of the Patient based on the re-positioning request issued from the PPRS.

· Note that for this profile, the final output of the transaction is the fact that the patient positioning has been successfully performed (including a correction or re-correction against an initial setup, when needed).

· After exiting this profile, typically the treatment session will continue with the Treatment Delivery Device (TDD) delivering the intended treatment.

3.1 Actors/ Transactions

<Actor/ Transaction diagram and Actor/ Transaction table. >

Figure X.1-1 shows the actors directly involved in the <Profile Name> Integration Profile and the relevant transactions between them.  Other actors that may be indirectly involved due to their participation in <other related profiles>, etc. are not necessarily shown.


[image: image6.wmf] 

­

 Transaction_3 [3]

 

­

 Transaction_4 [4]

 

 

Actor C

 

Transaction_1 [1] 

¯

 

¯

 Transaction _2 [2]

 

Actor A

 

Actor F

 

Actor D

 

Actor E

 

Actor B

 

Transaction_1 [1] 

­

 

­

 Transaction _2 [2]

 


Figure X.1-1.  <Profile Name> Actor Diagram

Table X.1-1 lists the transactions for each actor directly involved in the <Profile Name> Profile. In order to claim support of this Integration Profile, an implementation must perform the required transactions (labeled “R”). Transactions labeled “O” are optional.  A complete list of options defined by this Integration Profile and that implementations may choose to support is listed in Volume I, Section X.2.

Table X.1-1.  <Profile Name> Integration Profile - Actors and Transactions

	Actors
	Transactions 
	Optionality
	Section in Vol. 2

	TSM
	Transaction_1
	R
	Z.1

	
	Transaction_2
	R
	Z.2

	TDD
	Transaction_1
	R
	Z.1

	
	Transaction_2
	R
	Z.2

	PPD(s)
	Transaction_1
	R
	Z.1

	
	Transaction_2
	R
	Z.2

	
	Transaction_3
	R
	Z.3

	
	Transaction_4
	R
	Z.4

	PPVS
	Transaction_1
	R
	Z.1

	
	Transaction_2
	R
	Z.2

	
	Transaction_3
	R
	Z.3

	
	Transaction_4
	R
	Z.4

	PPRS
	Transaction_3
	R
	Z.3

	
	Transaction_4
	R
	Z.4


Note:
This is a note 

3.2 Discrete Positioning Workflow Integration Profile Options

<Table describing the options available for this integration profile and references to all sections that state requirements for compliance to the Option.>

Options that may be selected for this Integration Profile are listed in the table X.2-1 along with the Actors to which they apply.  Dependencies between options when applicable are specified in notes.

Table X.2-1 <Profile Name> - Actors and Options

	Actor
	Options
	Vol & Section

	Actor A
	No options defined 
	- -

	Actor F
	No options defined 
	- -

	Actor D/ Actor E
	No options defined 
	- -

	
	No options defined 
	- -

	Actor B (Actor C)
	No options defined 
	- -

	
	No options defined 
	- -


3.3 Discrete Positioning Workflow Process Flow

The process flow for the Discrete Positioning and Delivery Workflow Profile is shown in the following figures.

[image: image7.wmf]PPVS

TDD

TSM

Archive

Acquire Positioning Information

UPS Final Update 

[

RO

-

21

]

UPS Completed 

/ 

Cancelled 

[

RO

-

25

]

Store Position Acquisition Results to Archive 

[

RO

-

22

] (

Opt

)

Worklist Query for Positioning Acquisition 

[

RO

-

DPD

-

01

]

PPRS

Retrieve Device Position Information 

[

RO

-

DPD

-

02

]

Retrieve Workitem Input Object from TMS 

[

RO

-

20

]

TMS

UPS in Progress 

[

RO

-

19

]

User initiates acquisition

Version 

1

.

7

2011

-

05

-

13

Notify Device to start UPS 

[

RO

-

DPD

-

20

]

G

PPD

(

s

)

Retrieve Device Position Information 

[

RO

-

DPD

-

02

]


Figure 2.3‑1 IHE Discrete Positioning Workflow Integration Profile 1

[image: image8.wmf]PPVS

TDD

TSM

Archive

UPS Final Update 

[

RO

-

21

]

UPS Completed 

/ 

Cancelled 

[

RO

-

25

]

Perform Registration

UPS in Progress 

[

RO

-

19

]

Store Position Registration Results to Archive 

[

RO

-

23

]

Request RT Patient Position Correction 

[

RO

-

DPD

-

03

]

Calculated 

/ 

Decide on Patient Positioning

Store RT Patient Position Correction Instruction 

[

RO

-

DPD

-

04

]

Store RT Patient Position Correction Instruction 

[

RO

-

DPD

-

04

]

Retrieve Device Position Information 

[

RO

-

DPD

-

02

]

PPRS

Retrieve Positioning Acquisition Results 

[

RO

-

DPD

-

11

]

Retrieve Workitem Input Object from TMS 

[

RO

-

20

]

TMS

Worklist Query for Positioning Registration 

[

RO

-

DPD

-

12

]

Version 

1

.

7

2011

-

05

-

13

User initiates registration

Notify Device to start UPS 

[

RO

-

DPD

-

20

]

PPD

(

s

)


Figure 2.3‑2 IHE Discrete Positioning Workflow Integration Profile 2

[image: image9.wmf]PPD Aux

PPD Primary

TSM

Archive

Adjust Patient Position

Worklist Query for Repositioning 

[

RO

-

DPD

-

06

]

UPS Final Update 

[

RO

-

21

]

UPS Completed 

/ 

Cancelled 

[

RO

-

25

]

Store Reposition Results to Archive 

[

RO

-

DPD

-

05

]

Create new Positioning UPS

If the user decides to re

-

asses

the patient position

, 

the process restarts 

by creating new UPS as required

by RO

-

DPD

-

01

Retrieve RT Patient Position Correction Instruction 

[

RO

-

DPD

-

08

]

Adjust Patient Position

User initiates correction of patient position

Retrieve RT Patient Position Modification Instruction 

[

RO

-

DPD

-

14

] (

optional

)

Store RT Patient Position Modification Result 

[

RO

-

DPD

-

15

]

User authorizes and initiates Repositioning

User authorizes and initiates Repositioning

Version 

1

.

7

2011

-

05

-

13

Notify Device to start UPS 

[

RO

-

DPD

-

20

]

Notify Device to start UPS 

[

RO

-

DPD

-

20

] (

optional

)

Re

-

positioning succession

E

(

in case or Monitoring

)

F


Figure 2.3‑3 IHE Discrete Positioning Workflow Integration Profile 3
3.4 Discrete Positioning Workflow Security Considerations

TBD

4 Discrete Delivery Workflow Integration Profile

The Discrete Delivery Workflow Profile involves the treatment delivery performed after discrete patient positioning workflow. This profile is alternative to the 2 other profiles: Discrete Delivery and Monitoring Workflow Integration Profile, where monitoring is involved during the treatment delivery and the Integrated Positioning and Delivery Workflow Integration Profile, where a PDS actor does both imaging/monitoring and delivery.

· The Treatment Delivery System (TDD) retrieves the input objects (the treatment plan) from the Archive.

· The Treatment Delivery Device (TDD) delivers the intended treatment. Once the beam is on the Treat UPS status gets updated to being in progress. Once the beam is off the Treat UPS status gets updated to the final state.

· Finally, the TDD stores back the output objects (treatment records).

4.1 Actors/ Transactions

<Actor/ Transaction diagram and Actor/ Transaction table. >

Figure X.1-1 shows the actors directly involved in the <Profile Name> Integration Profile and the relevant transactions between them.  Other actors that may be indirectly involved due to their participation in <other related profiles>, etc. are not necessarily shown.


[image: image10.wmf] 

­

 Transaction_3 [3]

 

­

 Transaction_4 [4]

 

 

Actor C

 

Transaction_1 [1] 

¯

 

¯

 Transaction _2 [2]

 

Actor A

 

Actor F

 

Actor D

 

Actor E

 

Actor B

 

Transaction_1 [1] 

­

 

­

 Transaction _2 [2]

 


Figure X.1-1.  <Profile Name> Actor Diagram

Table X.1-1 lists the transactions for each actor directly involved in the <Profile Name> Profile. In order to claim support of this Integration Profile, an implementation must perform the required transactions (labeled “R”). Transactions labeled “O” are optional.  A complete list of options defined by this Integration Profile and that implementations may choose to support is listed in Volume I, Section X.2.

Table X.1-1.  <Profile Name> Integration Profile - Actors and Transactions

	Actors
	Transactions 
	Optionality
	Section in Vol. 2

	TSM
	Transaction_1
	R
	Z.1

	
	Transaction_2
	R
	Z.2

	TDD
	Transaction_1
	R
	Z.1

	
	Transaction_2
	R
	Z.2

	Check Performer
	Transaction_1
	R
	Z.1

	
	Transaction_2
	R
	Z.2

	
	Transaction_3
	R
	Z.3

	
	Transaction_4
	R
	Z.4


Note:
This is a note 

4.2 Discrete Delivery Workflow Integration Profile Options

<Table describing the options available for this integration profile and references to all sections that state requirements for compliance to the Option.>

Options that may be selected for this Integration Profile are listed in the table X.2-1 along with the Actors to which they apply.  Dependencies between options when applicable are specified in notes.

Table X.2-1 <Profile Name> - Actors and Options

	Actor
	Options
	Vol & Section

	Actor A
	No options defined 
	- -

	Actor F
	No options defined 
	- -

	Actor D/ Actor E
	No options defined 
	- -

	
	No options defined 
	- -

	Actor B (Actor C)
	No options defined 
	- -

	
	No options defined 
	- -


4.3 Discrete Delivery Workflow Process Flow

The process flow for the Discrete Monitoring Workflow Profile is shown in Figure Error! Reference source not found.-1.

[image: image11.wmf]Checker Performer

TDD

Archive

Deliver Radiation

Repeat for all Beams of current Treat UPS

TSM

Store Delivery Results to Archive 

[

RO

-

24

]

Notify on Beam Delivery Status Change 

[

RO

-

DPD

-

07

]

Notify on Beam Delivery Status Change 

[

RO

-

DPD

-

07

]

UPS Completed 

/ 

Cancelled 

[

RO

-

25

]

UPS Final Update 

[

RO

-

21

]

Worklist Query for Positioning and Delivery

[

RO

-

17

]

Version 

1

.

7

2011

-

05

-

13

Continue with subsequent Beam

Retrieve Worklist Input Objects from Archive 

[

RO

-

15

]

Request Validation of Treatment 

[

RO

-

XXX

] ((

TBD

, 

if it should be there

))

Notify Device to start UPS 

[

RO

-

DPD

-

20

]

UPS Progress Update for Treatment 

[

RO

-

26

]

TMS

Change Treat UPS Status  

[

RO

-

DPD

-

07

]

UPS Progress Update for Treatment 

[

RO

-

26

]

UPS in Progress 

[

RO

-

19

]

UPS Progress Update for Treatment 

[

RO

-

26

]


Figure 3.3‑1 IHE Discrete Delivery Workflow Integration Profile

4.4 Discrete Delivery Workflow Security Considerations

TBD

5 Discrete Delivery and Monitoring Integration Profile

The Discrete Monitoring and Delivery Workflow Profile involves the monitoring the patient position during treatment delivery, where the monitoring and delivery are performed by separate devices:

The overall workflow is as follows:

· The Patient Positioning Monitoring System (PPMS) typically acts as a Patient Positioning Verification System (PPVS) in the delivery preparation phase to bring the patient into the correct position prior to start the beam. Potentially this could be another PPVS as well. This initial positioning step therefore will be along the integrated delivery workflow profile or along the discrete positioning workflow profile.

· The Treatment Delivery System (TSM) pulls the worklist items and retrieves the input objects (the treatment plan) from the Archive.

· The PPMS subscribes to the Treatment Delivery Device (TSM) to request a notification, when the beam is switched on.

· The Treatment Delivery Device (TDD) informs the PPMS, which beam is activated and allows the PPMS to prepare for monitoring.

· The Treatment Delivery Device (TDD) then delivers the intended treatment. Once the beam is on, the Treat UPS status is updated. The PPMS get notified. It starts monitoring the patient position and determines, whether the patient is still in the tolerances provided for treatment.

· Once the beam is switched off, the Treat UPS status is updated. The PPMS gets notified about the deactivation or the beam. Once the PPMS has monitored all beams, which it is supposed to monitor, it may unsubscribe from the TDD.  

· The TDD stores back the output objects (treatment records).

When the patient position is determined as being out of tolerance by the PPMS, it is essential that the beam is switched off immediately. This operation is in a safety-critical real-time domain, which needs special considerations.

· The specific technology to switch off the beam in real-time is not specified in this Profile. This definition is left to an either eventually upcoming cross-vendor standard for a hardware interface specification to facilitate real-time beam hold signals – or as long as this standard is not available, a peer-to-peer interface definitions between a PPMS device and a TDD.

The actions taken to account for a patient position out-of-tolerance condition depends highly on the capabilities of the TDD and associated PPDs in respect to attended or automated compensation for the positional error. The following are typical scenarios:

1. The patient position is adjusted along the Discrete Positioning Workflow specified herein, which include re-acquisition of position reference data, registration and position correction.

2. The patient position is automatically corrected along a RT Patient Position Correction Instruction issued by the PPMS.

3. The user manually adjusts the patient position along information available on the PPMS.

4. The TDD and / or other actors could adapt beam parameters for actual delivery to account for the positional error.

5. The user terminates the treatment for this session.

The specification for the operations above is defined in the scenario diagrams in the next section.
5.1 Actors/ Transactions

<Actor/ Transaction diagram and Actor/ Transaction table. >

Figure X.1-1 shows the actors directly involved in the <Profile Name> Integration Profile and the relevant transactions between them.  Other actors that may be indirectly involved due to their participation in <other related profiles>, etc. are not necessarily shown.


[image: image12.wmf] 

­

 Transaction_3 [3]

 

­

 Transaction_4 [4]

 

 

Actor C

 

Transaction_1 [1] 

¯

 

¯

 Transaction _2 [2]

 

Actor A

 

Actor F

 

Actor D

 

Actor E

 

Actor B

 

Transaction_1 [1] 

­

 

­

 Transaction _2 [2]

 


Figure X.1-1.  <Profile Name> Actor Diagram

Table X.1-1 lists the transactions for each actor directly involved in the <Profile Name> Profile. In order to claim support of this Integration Profile, an implementation must perform the required transactions (labeled “R”). Transactions labeled “O” are optional.  A complete list of options defined by this Integration Profile and that implementations may choose to support is listed in Volume I, Section X.2.

Table X.1-1.  <Profile Name> Integration Profile - Actors and Transactions

	Actors
	Transactions 
	Optionality
	Section in Vol. 2

	TSM
	Transaction_1
	R
	Z.1

	
	Transaction_2
	R
	Z.2

	TDD
	Transaction_1
	R
	Z.1

	
	Transaction_2
	R
	Z.2

	PPMS
	Transaction_1
	R
	Z.1

	
	Transaction_2
	R
	Z.2

	TDD
	Transaction_1
	R
	Z.1

	
	Transaction_2
	R
	Z.2

	Check Performer
	Transaction_1
	R
	Z.1

	
	Transaction_2
	R
	Z.2

	
	Transaction_3
	R
	Z.3

	
	Transaction_4
	R
	Z.4


Note:
This is a note 

5.2 Discrete Delivery and Monitoring Integration Profile Options

<Table describing the options available for this integration profile and references to all sections that state requirements for compliance to the Option.>

Options that may be selected for this Integration Profile are listed in the table X.2-1 along with the Actors to which they apply.  Dependencies between options when applicable are specified in notes.

Table X.2-1 <Profile Name> - Actors and Options

	Actor
	Options
	Vol & Section

	Actor A
	No options defined 
	- -

	Actor F
	No options defined 
	- -

	Actor D/ Actor E
	No options defined 
	- -

	
	No options defined 
	- -

	Actor B (Actor C)
	No options defined 
	- -

	
	No options defined 
	- -


5.3 Discrete Delivery and Monitoring Process Flow

The process flow for the Discrete Monitoring Workflow Profile is shown in the following figures.

5.3.1 Overall Process Flow

[image: image13.wmf]PPMS

TDD

Archive

TSM

Retrieve Worklist Input Objects from Archive 

[

RO

-

15

]

Worklist Query for Positioning and Delivery

[

RO

-

17

]

Version 

1

.

7

2011

-

05

-

13

Notify Device to start UPS 

[

RO

-

DPD

-

20

]

Notify Device to start UPS 

[

RO

-

DPD

-

20

]

Worklist Query for Position Monitoring 

[

RO

-

DPD

-

13

]

Retrieve Worklist Input Objects from Archive 

[

RO

-

15

]

Subscribe for Treat UPS  

[

RO

-

DPD

-

09

]

Subscribe for Monitoring UPS  

[

RO

-

DPD

-

09

]


Figure 4.3‑1 IHE Discrete Monitoring Workflow Integration Profile 1

 [image: image14.wmf]PPMS

TDD

Archive

TSM

UPS in Progress 

[

RO

-

19

]

When PPMS detects

, 

that position is out of tolerance

, 

the PPMS will signal to the TDD that the beam shall be 

paused

. 

The real

-

time facilities to support this are 

described

 

in the text above

.

Notify on Beam Delivery Status Change 

[

RO

-

DPD

-

07

]

Store Position Registration Results to Archive 

[

RO

-

23

]

Store Monitoring Results 

[

RO

-

DPD

-

16

] (

Opt

)

Notify on active beam 

[

RO

-

DPD

-

10

]

Notify on active beam 

[

RO

-

DPD

-

10

]

Version 

1

.

7

2011

-

05

-

13

Deliver Radiation 

(

paused

)

Deliver Radiation

Various scenarios

, 

when 

out of tolerance

, 

are 

described in text above

.

Indicate Ready for Monitoring 

[

RO

-

DPD

-

19

]

TMS

UPS Progress Update for Treatment 

[

RO

-

26

]

UPS Progress Update for Treatment 

[

RO

-

26

]

B

C

D

Store Delivery Results to Archive 

[

RO

-

21

]

UPS Progress Update 

[

RO

-

xx

]

UPS Progress Update 

[

RO

-

xx

]

UPS in Progress 

[

RO

-

19

]

Notify on Beam Delivery Status Change  

[

RO

-

DPD

-

07

]


Figure 4.3‑2 IHE Discrete Monitoring Workflow Integration Profile 2

[image: image15.wmf]TSM

PPMS

TDD

Archive

UPS Completed 

/ 

Cancelled 

[

RO

-

25

]

Unsubscribe for Treat UPS  

[

RO

-

DPD

-

09

]

UPS Final Update 

[

RO

-

21

]

UPS Completed 

/ 

Cancelled 

[

RO

-

25

]

Version 

1

.

7

2011

-

05

-

13

Continue with subsequent Beam

C

Repeat for all Beams of current Treat UPS

UPS Final Update 

[

RO

-

21

]

UPS Completed 

/ 

Cancelled 

[

RO

-

25

]

Unsubscribe for Monitoring UPS


Figure 4.3‑3 IHE Discrete Monitoring Workflow Integration Profile 3

5.3.2 Interruption Handling by Re-Acquisition of Positioning Data

 [image: image16.wmf]TSM

TDD

G

B

User terminates the beam

UPS Final Update 

[

RO

-

21

]

UPS Completed 

/ 

Cancelled 

[

RO

-

25

]

Create Positioning Acquisition and Positioning Registration UPS 

[

RO

-

DPD

-

21

]

Version 

1

.

7

2011

-

05

-

13

TMS

Store Delivery Results to Archive 

[

RO

-

21

]

Create Treat UPS and Beam Delivery Instruction for continuation 

[

RO

-

DPD

-

22

]


Figure 4.3‑4 Interruption Handling by Re-Acquisition of Positioning Data

5.3.3 Interruption Handling by Automatic Re-positioning

 [image: image17.wmf]TSM

PPMS

TDD

B

Create new Positioning UPS

Create new RT Patient Position Correction Instruction

E

F

C

Notify Device to resume UPS

Version 

1

.

7

2011

-

05

-

13


Figure 4.3‑5 Interruption Handling by Automatic Re-positioning

5.3.4 Interruption Handling by Manual Re-positioning

 [image: image18.wmf]TSM

PPMS

TDD

B

E

F

C

Notify Device to resume UPS

User enters new position variables

Create new UPS and RT Patient Position Correction Instruction

Version 

1

.

7

2011

-

05

-

13


Figure 4.3‑6 Interruption Handling by Manual Re-positioning

5.3.5 Interruption Handling by Beam Parameter Adaptation

 [image: image19.wmf]TSM

TDD

B

C

Notify Device to resume UPS

User enters new position variables

Create new UPS and RT Patient Position Correction Instruction

Notify Device to start UPS 

[

RO

-

DPD

-

20

] (

optional

)

Worklist Query for Repositioning 

[

RO

-

DPD

-

06

]

Retrieve RT Patient Position Correction Instruction 

[

RO

-

DPD

-

08

]

Recalculate Beam Parameters to adjust for requested correction

UPS Final Update 

[

RO

-

21

]

UPS Completed 

/ 

Cancelled 

[

RO

-

25

]

Version 

1

.

7

2011

-

05

-

13


Figure 4.3‑7 Interruption Handling by Beam Parameter Adaptation

5.3.6 Interruption Handling by Treatment Termination

 [image: image20.wmf]TDD

B

C

User terminates the beam

Version 

1

.

7

2011

-

05

-

13


Figure 4.3‑8 Interruption Handling by Treatment Termination

5.4 Discrete Delivery and Monitoring Security Considerations

TBD

<Appendix A> Actor Summary Definitions
Patient Positioning Verification System (PPVS) – This actor represents a system responsible for acquiring data representing the patient position before the beam delivery. The PPVS fulfils the role of a UPS-Pull ‘Pull Performer’ SCU as described in DICOM Supplement 96 Part 17 Table Z.1-1.

Patient Positioning Registration System (PPRS) – This actor represents a system responsible for registering data acquired, which represent the patient position, against reference data containing information about the intended patient position, before the beam delivery. The PPRS will expose the difference between intended and actual position (or the requested correction), but will not be involved in the re-positioning process itself. The PPRS fulfils the role of a UPS-Pull ‘Pull Performer’ SCU as described in DICOM Supplement 96 Part 17 Table Z.1-1.

Patient Positioning Monitoring System (PPMS) – This actor represents a system responsible for determining patient positioning during the beam delivery, determining the patient position in real-time and comparing it against reference data. When this actor determines, that the patient position goes out of a certain tolerance, it will issue a signal to the TDD to request the beam to be paused. The PPMS fulfils the role of a UPS-Pull ‘Pull Performer’ SCU as described in DICOM Supplement 96 Part 17 Table Z.1-1.

Treatment Session Manager (TSM) – An actor managing the different activities during a treatment session. It serves as a local Worklist SCP to allow the different devices involved in the execution of the treatment session to pull their UPS and update the UPS status. This actor is especially important for the discrete positioning and delivery workflow, but may eventually appear in others context too.

Patient Positioning Device (PPD) – An actor performing part or all of the patient positioning as a positioning device.  Note, that the TDD actor normally has capability for patient positioning as well. Therefore this device normally serves the role of a primary PPD. However, this device may only cover some axis respectively not perform the whole range of patient positioning. In this case, eventually secondary PPD actor(s) maybe present as a separate positioning device additional to the TDD as primary PDD. 

Treatment Delivery Device (TDD) – The TDD has both the facilities of delivering the beam and positioning the beam (not the patient) in the room coordinate system. This profile makes the explicit assumption, that those 2 systems are combined in one actor in respect to the others of the profile.
<Appendix B> Transaction Summary Definitions

RO-DPD-01: Worklist Query for Positioning Acquisition

A PPVS queries the worklist server to get the steps for positioning.

RO-DPD-02: Retrieve Device Position Information

A PPVS or PPMS requests the actual device parameters from the treatment delivery device (TDD) and patient positioning device(s) (PPD(s))..

RO-DPD-03: Request RT Patient Position Correction

A PPVS or PPMS request a correction of the patient position to be performed prior to treatment.

RO-DPD-04: Store RT Patient Position Correction Instruction

A PPVS stores the RT position correction instruction to the TSM respectively the Archive specifying the request correction of the patient position.

RO-DPD-05: Store RT Repositioning Results to Archive

The TDD reports the re-positioning to the Archive with a RT Patient Positioning Modification Result.

RO-DPD-06: Worklist Query for Repositioning

The TDD queries the worklist to get the request for modification of the patient position and retrieves the RT Patient Positioning Modification Instruction to perform the re-positioning.

RO-DPD-07: Notify on Beam Delivery Status Change
The TDD notifies the TSM and the PPMS to about changes in the beam delivery status.
RO-DPD-08: Retrieve RT Patient Position Correction Instruction

A TDD retrieves the RT position correction instruction from the TSM, which specifies the request correction of the patient position.

RO-DPD-09: Subscribe/Unsubscribe to Treat UPS Status

The PPMS subscribes / unsubscribe to changes of the Treat UPS to get the status of beam on / off.

RO-DPD-10: Notify on active beam

The TDD notifies the PPMS on activation / decactivation of the beam.

RO-DPD-11: Retrieve Positioning Acquisition Results

A PPRS retrieves the position acquisition result from the Archive to be able to calculate the difference between intended patient position and actual patient position.

RO-DPD-12: Worklist Query for Positioning Registration

A PPRS queries the worklist server to get the steps for registration.

RO-DPD-13: Worklist Query for Position Monitoring

PPMS queries the worklist server to get the steps for Positioning monitoring.

RO-DPD-14: Retrieve RT Patient Position Modification Instruction

A PPD retrieves a RT Patient Position Modification Instruction, which contains the portion of the position modification, which is left to the PPD.

RO-DPD-15: Store RT Patient Position Modification Result

A PPD stores back the performed positioning modification result.

RO-DPD-16: Store Monitoring Results to Archive

A PPMS stores back the result of monitoring.

RO-DPD-17: Collective UPS Final Update

Once the treatment session is finished, the TSM updates the UPS states in the TMS with the final content information.

RO-DPD-18: Collective UPS Completed / Cancelled

Once the treatment session is finished, the TSM updates the UPS states in the TMS with the final status.

RO-DPD-19: Indicate Ready for Monitoring

The PPMS indications to the TDD, that it is ready for monitoring the patient position.

RO-DPD-20: Notify Device to start UPS

Some actor (in most cases the TSM) notifies a device, that it the treatment session process is in a state, that the next UPS can be performed.

RO-DPD-21: Create Positioning Acquisition and Positioning Registration UPS

When an interruption of the treatment occurred (especially in the context of monitoring), the TMS will create new UPS to initiate a re-acquisition of positioning information and to perform a registration to re-validate the patient position
RO-DPD-22: Create Treat UPS and Beam Delivery Instruction for Continuation

When the treatment is resumed after an interruption with the session, the TSM creates a new UPS and a Beam Delivery Request IOD to instruct the TDD on the continuation of the treatment.
Glossary

Add the following terms to the Glossary:

<any glossary additions associated with the profile draft go here>

Volume 2 - Transactions

Add section 3.Y 

5.5 RO-DPD-01: Worklist Query for Positioning Acquisition

5.5.1 Scope

In the Worklist Query for Position Acquisition transaction, the performing device requests and receives a patient positioning worklist from a worklist server. In the integrated delivery workflow, the performing device is the Positioning and Delivery System and the worklist server is the Treatment Management System. In the discrete positioning profile, the performing device is the Patient Positioning Verification System and the worklist server is the Treatment Session Manager.

5.5.2 Use Case Roles


[image: image21.emf]Worklist Query for Positioning

Treatment Management System (TMS)

or

Treatment Session Manager (TSM)

Positioning and Delivery System (PDS)

or

Patient Positioning Verification System (PPVS)


Actor: Positioning and Delivery System or Patient Positioning Verification System (‘Performing Device’)
Role: Queries a Treatment Management System or Treatment Session Manager and receives a scheduled patient positioning worklist.

Actor: Treatment Management System or Treatment Session Manager ("Worklist Server")

Role: Responds to a worklist query and sends a scheduled patient positioning worklist to a Positioning and Delivery System or Patient Positioning Verification System.

5.5.3 Referenced Standards

DICOM 2010 PS 3.4: Annex UUU Unified Procedure Step Service
5.5.4 Interaction Diagram


[image: image22.emf]PDS or PPVS TMS or TSM

Query Scheduled UPS Worklist

Receive Scheduled UPS Worklist


5.5.4.1 Query Scheduled UWPS Worklist Message

This is the worklist query message sent to the worklist server.

5.5.4.1.1 Trigger Events

The user of the performing device, in order to position a patient prior to treatment delivery, requests that the worklist server to send a scheduled patient positioning worklist.

5.5.4.1.2 Message Semantics

For the Worklist Query for Position Acquisition, exactly one Unified Procedure Steps (UPS C-FIND responses in the ‘pending’ state) shall be returned for each positioning verification:

· Scheduled Workitem Code Sequence (0040,4018) Code Value shall be in the range 121702-121711 (acquisition) and Coding Scheme Designator shall be ‘DCM’.
The Input Information Sequence (0040,4021) shall contain a RT Patient Position Setup Reference Acquisition Instruction IOD
If the Code Value is in the range 121702-121706, one or more RT Image SOP Instance references shall be supplied in the Input Information Sequence (0040,4021).  If the Code Value is in the range 121707-121708, one CT Image Series reference shall be supplied in the Input Information Sequence (0040,4021)
The specified location of the SOP Instances shall be an Archive.

For father details, see Appendix A.

5.5.4.1.3 Expected Actions

The performing device reads the instruction IODs for positioning, prepares acquisition of the requested positioning data along the instructions and waits for the user to initiate acquisition..
5.6 RO-DPD-02: Retrieve Patient Positioning Device Actuals

5.6.1 Scope

The Patient Positioning Verification System or Patient Positioning Monitoring System requests and receives the actual patient position parameters from the Treatment Delivery Device and the Patient Positioning Device(s). This is needed for both to record the actual position in the object representing the acquired data (e.g. images etc.) and to determine the new position in case of position corrections needed.

5.6.2 Use Case Roles

 [image: image23.emf]Retrieve Patient Positioning Device 

Actuals

Treatment Delivery Device (TDD)

or

PatientDelivery Device (PDD)

Patient Positioning System (PPVS)

or

Patient Positioning System (PPMS)


Actor: Patient Positioning Verification System or Patient Positioning Monitoring System (‘Performing Device’)
Role: .
Actor: Treatment Delivery Device or Patient Positioning device
Role: Responds to a request to retrieve the actual patient position information.

5.6.3 Referenced Standards

TBD

5.6.4 Interaction Diagram

[image: image24.wmf]PPVS

, 

PPMS

TDD

Request actual patient position information

Receive actual patient position information


5.6.4.1 Request actual patient position information

This is the request to get the actual patient position information from the treatment delivery device.

5.6.4.1.1 Trigger Events

Before the performing device is acquiring images, marker position or similar, resp. calculating patient position corrections, it may retrieve the actual patient position from the treatment delivery device.

5.6.4.1.2 Message Semantics

((TBD))

5.6.4.1.3 Expected Actions

The performing device uses the actual information for determining the necessary corrections for the patient position.

5.7 RO-DPD-03:  Request RT Patient Position Correction

5.7.1 Scope

The Patient Positioning Registration System stores the instruction to modify the patient position to the Treatment Session Manager and creates a UPS for performing the re-positioning.

5.7.2 Use Case Roles


[image: image25.emf]Create UPS for Re-positioning

Treatment Session 

Manager (TSM)

Patient Positioning 

Registration System 

(PPRS)


Actor: Patient Positioning Registration System ("Performing Device")

Role: Store Position Modification Instruction
Creates UPS to perform re-positioning

Actor: Treatment Session Manager

Role: Receives Position Modification Instruction and UPS

5.7.3 Referenced standards

DICOM 2008 PS3.4: Storage Service Class.

DICOM 2010 PS 3.4: Annex UUU Unified Procedure Step Service
DICOM Supplement 147: Second Generation Radiotherapy

5.7.4 Interaction Diagram


[image: image26.emf]TSM PPRS

N-CREATE (UPS)


5.7.4.1 Create UPS for RT Patient Position Correction Instruction

5.7.4.1.1 Trigger Events

When the user has initiated the request for re-positioning and the PPRS has sent the RT Position Modification Instruction, the PPRS creates a UPS for Patient Positioning.
5.7.4.1.2 Message Semantics

The Scheduled Workitem Code Sequence (0040,4018) Code Value shall be equal to ‘121722’ (‘RT Patient Position Adjustment’) and Coding Scheme Designator shall be equal to ‘DCM’.

The Input Information SOP Instance sequence shall contain the reference to the RT Patient Position Correction Instruction IOD, which has been stored by the previously by PPRS to the Treatment Management System and Treatment Session Manager.

For father details, see Appendix A.

5.7.4.1.3 Expected Actions

The TSM will create the UPS for Re-positioning. This will be handled as an N-ACTION as defined by the DIMSE Group UPS Push.
5.8 RO-DPD-04:  Store RT Patient Position Correction Instruction
5.8.1 Scope

The Positioning and Delivery System (in the Integrated Delivery Profile) or the Patient Positioning Registration System (in the Discrete Positioning and Delivery Profile) stores the requested correction of the patient position to the Treatment Management System respectively the Treatment Session Manager.

5.8.2 Use Case Roles


[image: image27.emf]Store RT Position Modification 

Instruction

Positioning and Delivery System (PDS)

or

Patient Positioning System (PPRS)

Treatment Management System (TMS)

or

Treatment Session Manager (TSM)


Actor: Positioning and Delivery System or Patient Positioning Registration System ("Performing Device")
Role: Store RT Position Modification IOD

Actor: Treatment Management System or Treatment Session Manager

Role: Receives RT Position Modification IOD

5.8.3 Referenced standards

DICOM 2008 PS3.4: Storage Service Class.

DICOM Supplement 147: Second Generation Radiotherapy

5.8.4 Interaction Diagram


[image: image28.emf]TMS or TSM

PDS or PPRS

C-STORE (RT Patient Position Modification Instruction)


5.8.4.1 Store RT Patient Position Correction Instruction

5.8.4.1.1 Trigger Events

The user of the performing device has performed a registration and decided, whether a re-positioning is required or not. If a re-positioning is required, he determines the necessary position correction. To achieve the optimal patient position for treatment.

The user initiates the performing device to send the RT Patient Position Modification Instruction IOD.

The performing device uses the DICOM C-STORE message to transfer the RT Patient Position Correction Instruction to the TSM (in case of discrete positioning and delivery profile) or Archive (in all profiles)

5.8.4.1.2 Message Semantics

The message semantics are defined by the DICOM Object Storage SOP Classes.

A C-STORE Request shall be sent from the Performing Device to the TSM and/or Treatment Management System. A single object shall be stored, with one of the SOP Classes denoted in the Table 3.20-1:

Table 3.20-1: Permitted SOP Class Support for Performing Device (SCU)

	SOP Class Name
	SOP Class UID

	RT Patient Position Correction Instruction
	


5.8.4.1.3 Expected Actions

The TMS will store the result for post-treatment processing and review.

The TSM will store the RT Position Modification Instruction as received, to make it available for the TDD later on.

5.9 RO-DPD-05: Store RT Repositioning Results to Archive
5.9.1 Scope

The Treatment Delivery Device stores the result of the re-positioning to the Archive.

5.9.2 Use Case Roles


[image: image29.emf]Store Repositioning Results

Archive

Treatment Delivery 

Device (TDD)


Actor: Treatment Delivery Device
Role: Stores the output of the re-positioning operation to the Archive.

Actor: Archive

Role: Responds to a C-STORE request and stores the transmitted objects.

5.9.3 Referenced standards

DICOM 2008 PS3.4: Storage Service Class.

DICOM Supplement 147: Second Generation Radiotherapy

5.9.4 Interaction Diagram


[image: image30.emf]Archive TDD

C-STORE (Repositioning Results)


5.9.4.1 
Storage of the RT Repositioning Results

5.9.4.1.1 Trigger Events

When the TDD has performed the re-positioning, it stores the re-positioning result to the Archive.

5.9.4.1.2  Message Semantics

The message semantics are defined by the DICOM Object Storage SOP Classes.

A C-STORE Request shall be sent from the Performing Device to the Archive

Table 3.20-1: Permitted SOP Class Support for Performing Device (SCU)

	SOP Class Name
	SOP Class UID

	RT  Patient Position Modification Result
	


5.9.4.1.3 Expected Actions

The TDD will recognize the UPS requesting a re-positioning.

5.10 RO-DPD-06: Worklist Query for Repositioning
5.10.1 Scope

The Treatment Delivery Device retrieves the UPS for re-positioning and the Repositioning Instruction from the Treatment Session Manager.

5.10.2 Use Case Roles

[image: image31.wmf]Worklist Query for Repositioning

Treatment Session Manager 

(

TSM

)

Treatment Delivery Device 

(

TDD

)


Actor: Treatment Delivery Device

Role: Receives UPS and Position Modification Instruction

Actor: Treatment Session Manager

Role: Store Position Modification Instruction
Creates UPS to perform re-positioning

5.10.3 Referenced standards

DICOM 2010 PS 3.4: Annex UUU Unified Procedure Step Service
DICOM Supplement 147: Second Generation Radiotherapy

5.10.4 Interaction Diagram

[image: image32.wmf]TSM

TDD

Query Scheduled UPS Worklist

Receive Scheduled UPS Worklist


5.10.4.1 Retrieve UPS for Repositioning

5.10.4.1.1 Trigger Events

The trigger to initiate the TDD to query for re-positioning may be an out-of-band protocol between the TSM and the TDD or a user interaction.

The TDD retrieves the UPS for Re-positioning from the TSM.
5.10.4.1.2 Message Semantics

For the Worklist Query for Position Acquisition, exactly one Unified Procedure Steps (UPS C-FIND responses in the ‘pending’ state) shall be returned for the re-positioning request.

5.10.4.1.3 Expected Actions

The TDD will recognize the UPS requesting a re-positioning of the patient.

5.11 RO-DPD-07: Notify on Beam Delivery Status Change
5.11.1 Scope

The Treatment Delivery Device changes the Treat UPS status and notifies the Treatment Session Manager or the subscribed Patient Positioning Monitoring System.

5.11.2 Use Case Roles

[image: image33.wmf]Notify on status change of

Treat UPS

Treatment delivery Device 

(

TDD

)

Treatment Session Manager 

(

TSM

)

or

Patient Position Monitoring System 

(

PPMS

)


Actor: Treatment Delivery System
Role: Sends an N-EVENT-REPORT, when the Treat UPS is started and when it is finished.

Actor: Treatment Session Manager or Patient Positioning Monitoring System

Role: Receives notification on status change of Treat UPS

5.11.3 Referenced standards

DICOM 2010 PS 3.4: Annex UUU Unified Procedure Step Service
5.11.4 Interaction Diagram

[image: image34.wmf]TDD

TSM or PPMS

N

-

EVENT

_

REPORT


5.11.4.1 Send status event notification to subscribed PPMS

The TDD will send the status changes to the subscribed TSM and PPMS. This will be handled as an N-EVENT-REPORT as defined by the DIMSE Group UPS EVENT.

5.11.4.1.1 Trigger Events

The TSM and/or PPMS will be notified that the beam delivery has been started or finished.

Note that this is not a notification fulfilling any deterministic real-time requirements. However, the notification should be as near as possible to beam on.

5.11.4.1.2 Message Semantics

5.11.4.1.3 Expected Actions

To ensure, that the notification is received priori to beam on, the TDD shall not enable beam before getting the N-EVENT-REPORT response confirming the reception of the N-EVENT-REPORT.
The PPMS will start monitoring, if applicable.

5.12 RO-DPD-08: Retrieve RT Patient Position Correction Instruction
5.12.1 Scope

The Treatment Delivery Device retrieves the requested correction of the patient position from Treatment Session Manager.

5.12.2 Use Case Roles


[image: image35.emf]Retrieve RT Position Modification 

Instruction

Treatment Delivery Device (TDD) Treatment Session Manager (TSM)


Actor: Treatment Delivery Device ("Performing Device")

Role: Retrieves RT Patient Position Modification Instruction IOD

Actor: Treatment Session Manager

Role: Provides RT Patient Position Modification Instruction IOD

5.12.3 Referenced standards

DICOM 2008 PS3.4: Storage Service Class.

DICOM Supplement 147: Second Generation Radiotherapy

5.12.4 Interaction Diagram

[image: image36.wmf]TSM

TDD

C

-

STORE 

(

RT Patient Position Modification Instruction

)

C

-

MOVE 

(

RT Patient Position Modification Instruction

)


5.12.4.1 Retrieve RT Patient Position Correction Instruction

The Retrieve (Study Root – MOVE) SOP Class shall be supported, with Series support. . Implementations shall support a mode of operation in which specific SOP Instances (rather than entire studies) are retrieved from the Treatment Management System using the Study Root – MOVE SOP Class. Refer to DICOM 2007 PS 3.4, Annex C, for detailed descriptive semantics.

A Performing Device shall be capable of issuing Study-Root C-MOVE for all relevant SOP Instances that are specified in the Input Information Sequence.

5.12.4.1.1 Trigger Events

The user of the PPVS has performed a registration and decided, whether a re-positioning is required or not. A UPS has been created in the TSM for re-positioning.

Based on this UPS, the TDD retrieves the RT Position Modification Instruction, containing the parameters how to re-position the patient.

5.12.4.1.2 Message Semantics

The message semantics are defined by the DICOM Object Storage SOP Classes.

Table 3.20-1: Permitted SOP Class Support for Performing Device (SCU)

	SOP Class Name
	SOP Class UID

	RT  Patient Position Correction Instruction Storage
	


5.12.4.1.3 Expected Actions

Based in the received RT  Patient Position Modification Instruction, the TDD shall propose the requested position correction to the user. The user than will invoke the necessary operations on the TDD to realize the correction (e.g. to move the patient support device accordingly). Note however, that the final control of what correction is applied in which amount is with the user operating the TDD.
5.13 RO-DPD-09: Subscribe / Unsubscribe to Treat UPS
5.13.1 Scope

The .Patient Positioning Monitoring System subscribes / unsubscribe to changes of the UPS Treat step to get notified about beam on / beam off events.

5.13.2 Use Case Roles

[image: image37.wmf]Subscribe 

/ 

Unsubscribe to UPS 

Treat Event

Treatment delivery 

Device 

(

TDD

)

Patient Position 

Monitoring System 

(

PPMS

)


Actor: Patient Positioning Monitoring System "Performing Device")

Role: Issues an N-ACTION to subscribe or unsubscribe the Treat UPS changes

Actor: Treatment Delivery System
Role: Receives the subscription.

5.13.3 Referenced standards

DICOM 2010 PS 3.4: Annex UUU Unified Procedure Step Service
5.13.4 Interaction Diagram

[image: image38.wmf]TDD

PPMS

N

-

ACTION


5.13.4.1 
Subscribe to UPS

5.13.4.1.1 Trigger Events

When the PPMS is activated (typically after this system has acted as a PPVS to determine the patient's initial position), it subscribes to the TDD to get notification when the Treat UPS is started, which corresponds with turning the beam on.

Once the delivery is finished and the beam is terminated normally or under abnormal conditions, the PPMS will get notified again by the TDD about this event.

5.13.4.1.2  Message Semantics

TBD

5.13.4.1.3 Expected Actions

The TDD will read the subscription. It will start notifying the PPMS on beam on/off events. This will be handled as an N-ACTION as defined by the DIMSE Group UPS Watch.
5.13.4.2 
Unsubscribe to UPS

5.13.4.2.1 Trigger Events

When the PPMS has finished monitoring for this patient's treatment session or has been switched off, it may unsubscribe to the status change event of the Treat UPS.

5.13.4.2.2  Message Semantics

TBD

5.13.4.2.3 Expected Actions

The TDD will read the unsubscribe request and will stop notifying the PPMS on beam on/off events.. This will be handled as an N-ACTION as defined by the DIMSE Group UPS Watch.

5.14 RO-DPD-10: Notify on active beam
5.14.1 Scope

The Treatment Delivery Device indicates to the Patient Positioning Monitoring System that a beam has been activated for treatment. This allows the Patient Positioning Monitoring System to prepare for monitoring. Once the beam has been delivered, this notification is used to indicate the de-activation.

5.14.2 Use Case Roles

Actor: Treatment Delivery Device (Performing Device)

Role: Issues the notification

Actor: Patient Positioning Monitoring System

Role: Receives the Beam activation / deactivation status

5.14.3 Referenced standards

DICOM 2008 PS3.4: Storage Service Class.

5.14.4 Interaction Diagram

[image: image39.wmf]Notify on active beam

Treatment delivery 

Device 

(

TDD

)

Patient Position 

Monitoring System 

(

PPMS

)


5.14.4.1 Notify on activation of beam

5.14.4.1.1 Trigger Events

Once a beams is activated, i.e. selected for treatment by any means at the TDD (User interaction or auto-sequencing), this event report is sent. At this point, radiation is not yet switched on.

5.14.4.1.2 Message Semantics

[image: image40.wmf]TDD

PPMS

N

-

EVENT

-

REPORT


5.14.4.1.3 Expected Actions

On receiving the activation, the PPMS will take necessary action to prepare the monitoring process.

5.15 RO-DPD-11: Retrieve Positioning Acquisition Results
5.15.1 Scope

The Patient Positioning Registration System retrieves the IOD containing the positioning acquisition results from the archive to be able to perform the registration.

5.15.2 Use Case Roles

Actor: Patient Positioning Registration System (Performing Device)

Role: Invokes a query/retrieve to the positioning acquisition results

Actor: Archive

Role: provides the acquisition results

5.15.3 Referenced standards

DICOM 2008 PS3.4: Storage Service Class.

5.15.4 Interaction Diagram


[image: image41.emf]Retrieve Positioning Acquisition 

Results

Archive

Patient Positioning 

Registration System 

(PPRS)


5.15.4.1 Retrieve Positioning Acquisition Results

5.15.4.1.1 Trigger Events

Once the PPRS has retrieves the UPS requesting the registration, it will retrieve the acquisition result IOD as referenced in eth UPS input sequence.

5.15.4.1.2 Message Semantics


[image: image42.emf]Archive PPRS

Retrieve Objects (C-MOVE)

Store Objects (C-STORE)


5.15.4.1.3 Expected Actions

Using the acquired position data, the PPRS initiates the registration of those data against reference data.

5.16 RO-DPD-12: Worklist Query for Positioning Registration

5.16.1 Scope

In the Worklist Query for Registration transaction, a Positioning and Delivery System (in the Integrated Delivery Profile) or the Patient Positioning Registration System (in the Discrete Positioning and Delivery Profile) requests and receives a patient positioning worklist from a Treatment Management System in the integrated delivery profile or the Treatment Management System or Treatment Session Manager in the discrete positioning and delivery profile.

5.16.2 Use Case Roles


[image: image43.emf]Worklist Query for Registration

Treatment Management System (TMS)

or

Treatment Session Manager (TSM)

Positioning and Delivery System (PDS)

or

Patient Positioning Registration System (PPRS)


Actor: Positioning and Delivery System or Patient Positioning Registration System (‘Performing Device’)
Role: Queries a Treatment Management System or Treatment Session Manager and receives a scheduled patient positioning worklist.

Actor: Treatment Management System or Treatment Session Manager ('Worklist Server')

Role: Responds to a worklist query and sends a scheduled patient positioning worklist to a Patient Positioning Registration System or Positioning and Delivery System.

5.16.3 Referenced Standards

DICOM 2010 PS 3.4: Annex UUU Unified Procedure Step Service
DICOM Supplement 147: Second Generation Radiotherapy

5.16.4 Interaction Diagram


[image: image44.emf]PDS or PPRS TMS or TSM

Query Scheduled UPS Worklist

Receive Scheduled UPS Worklist


5.16.4.1 Query Scheduled UWPS Worklist Message

This is the worklist query message sent to the Treatment Management System or Treatment Session Manager.

5.16.4.1.1 Trigger Events

The user of the performing device, in order to register the actual patient position against the reference positions prior to treatment delivery, requests that the worklist server sends a scheduled patient positioning registration worklist.

5.16.4.1.2 Message Semantics

For the Worklist Query for Positioning Registration, exactly one Unified Procedure Steps (UPS C-FIND responses in the ‘pending’ state) shall be returned for each positioning verification: 

· Scheduled Workitem Code Sequence (0040,4018) Code Value shall be in the range 121712-121721 (registration) and Coding Scheme Designator shall be ‘DCM’.
The Input Information Sequence (0040,4021) shall contain references to the position reference and the acquired positioning IODs and may contain an RT Registration Instruction.

For further details, see Appendix A.

5.16.4.1.3 Expected Actions

The performing device reads the instruction IODs for registrations, prepares registration of the acquired positioning data against the reference positioning data.

5.17 RO-DPD-13: Worklist Query for Position Monitoring

5.17.1 Scope

In the Worklist Query for Registration transaction, a Positioning and Delivery System (in the Integrated Delivery Profile) or the Patient Positioning Monitoring System (in the Discrete Positioning and Delivery Profile) requests and receives a patient positioning worklist from a Treatment Management System in the integrated delivery workflows or the Treatment Management System or Treatment Session Manager in the discrete positioning and delivery workflow.

5.17.2 Use Case Roles


[image: image45.emf]Worklist Query for Monitoring

Treatment Management System (TMS)

or

Treatment Session Manager (TSM)

Positioning and Delivery System (PDS)

or

Patient Positioning Monitoring System (PPMS)


Actor: Positioning and Delivery System or Patient Positioning Monitoring System (‘Performing Device’)
Role: Queries a Treatment Management System or Treatment Session Manager and receives a scheduled patient positioning worklist.

Actor: Treatment Management System or Treatment Session Manager (Worklist Server)

Role: Responds to a worklist query and sends a scheduled patient positioning worklist to a Patient Positioning Registration System or Positioning and Delivery System.
5.17.3 Referenced Standards

DICOM 2010 PS 3.4: Annex UUU Unified Procedure Step Service
DICOM Supplement 147: Second Generation Radiotherapy

5.17.4 Interaction Diagram


[image: image46.emf]PDS or PPMS TMS or TSM

Query Scheduled UPS Worklist

Receive Scheduled UPS Worklist


5.17.4.1 Query Scheduled UWPS Worklist Message

This is the worklist query message sent to the Treatment Management System or Treatment Session Manager.

5.17.4.1.1 Trigger Events

The user of the performing device, in order to monitor a patient position during treatment delivery, requests that the worklist server send a scheduled patient positioning worklist.

5.17.4.1.2 Message Semantics

For the Worklist Query for Monitoring, exactly two Unified Procedure Steps (UPS C-FIND responses in the ‘pending’ state) shall be returned for each monitoring activity.

· Response 1: Scheduled Workitem Code Sequence (0040,4018) Code Value shall be in the range ((TBD)) (monitoring) and Coding Scheme Designator shall be ‘DCM’.
The Input Information Sequence (0040,4021) may contain references to the position reference and the acquired positioning IODs and may contain a RT Patient Positioning Monitoring Instruction IOD

· Response 2: Scheduled Workitem Code Sequence (0040,4018) Code Value shall be in the range 121712-121721 (registration) and Coding Scheme Designator shall be ‘DCM’. 

For further details, see Appendix A.

5.17.4.1.3 Expected Actions

The PPMS reads the instruction IODs for monitoring, prepares monitoring the patient position along the instructions and waits TDD to initiate treatment.
5.18 RO-DPD-14: Retrieve RT Patient Position Modification Instruction

5.18.1 Scope

A Patient Positioning Device retrieves from the Treatment Delivery Device the remaining part of a patient position modification, which it is supposed to perform.

5.18.2 Use Case Roles


[image: image47.emf]Retrieve RT Position Modification 

Instruction

Patient Positioning Device (PPD) Treatment Delivery Device (TDD)


Actor: Patient Positioning Device
Role: Retrieves partial RT Patient Position Modification Instruction IOD

Actor: Treatment Delivery Device

Role: Provides partial RT Patient Position Modification Instruction IOD

5.18.3 Referenced Standards

DICOM 2010 PS 3.4: Annex UUU Unified Procedure Step Service
DICOM Supplement 147: Second Generation Radiotherapy

5.18.4 Interaction Diagram


[image: image48.emf]TDD

PPD

C-STORE (RT Patient Position Modification Instruction)

C-MOVE (RT Patient Position Modification Instruction)


5.18.4.1 Retrieve RT Patient Position Modification Instruction

5.18.4.1.1 Trigger Events

The user of the PPD, in order to re-position the patient before treatment delivery, requests, that the PPD performs the re-positioning.

5.18.4.1.2 Message Semantics

The PPD queries the TDD for a RT Patient Position Modification Instruction. The TDD will provide exactly one object.

If no action is expected from the PPD, the query will return with an error code of ((TBD)).

5.18.4.1.3 Expected Actions

On receiving the query from the PPD, the TDD constructs a RT Patient Position Modification Instruction IOD, which contains the remaining portion of the repositioning for the current patient position. The content of the transformation is calculated along what the TDD has already performed on re-positioning.

If there is no remaining re-positioning to be done, the query will return an error code as defined above/

5.19 RO-DPD-15: Store RT Patient Position Modification Result

5.19.1 Scope

A Patient Positioning Device stores the patient reposition results, which it has performed to the Treatment Delivery Device. This will allow the Treatment Delivery Device to merge this information into its T Patient Position Modification Result IOD to represent the total re-positioning.

5.19.2 Use Case Roles


[image: image49.emf]Store Repositioning Results

Treatment Delivery 

Device (TDD)

Patient Positioning 

Device (TDD)


Actor: Patient Positioning Device
Role: Stores the RT Patient Position Modification Result IOD

Actor: Treatment Delivery Device

Role: Receives the  RT Patient Position Modification Result IOD

5.19.3 Referenced Standards

DICOM 2010 PS 3.4: Annex UUU Unified Procedure Step Service
DICOM Supplement 147: Second Generation Radiotherapy

5.19.4 Interaction Diagram


[image: image50.emf]TDD PPD

C-STORE (Repositioning Results)


5.19.4.1 Store RT Patient Position Modification Result

5.19.4.1.1 Trigger Events

When the PPD has finished the re-positioning, it stores back the re-positioning result to the TDD.

5.19.4.1.2 Message Semantics

The message semantics are defined by the DICOM Object Storage SOP Classes.

A C-STORE Request shall be sent from the Performing Device to the TDD.

5.19.4.1.3 Expected Actions

On receiving the RT Patient Position Modification Result from the PPD, the TDD integrates the data from the PPD into its overall RT Patient Position Modification Result IOD, which it will send finally back to the Treatment Management System 

5.20 RO-DPD-16: Store Monitoring Results to Archive

5.20.1 Scope

A Positioning and Delivery System (in the Integrated Delivery Profile) or a Patient Positioning Monitoring System (in the Discrete Positioning and Delivery Profile) stores the monitoring results to the Archive.

5.20.2 Use Case Roles


[image: image51.emf]Store Monitoring Results

Archive

Positioning and Delivery System (PDS)

or

Patient Positioning Monitoring System (PPMS)


Actor: Positioning and Delivery System or Patient Positioning Monitoring System ("Performing Device")
Role: Stores the RT Patient Positioning Monitoring Result IOD

Actor: Archive

Role: Receives RT Patient Positioning Monitoring Result IOD

5.20.3 Referenced Standards

DICOM 2010 PS 3.4: Annex UUU Unified Procedure Step Service
DICOM Supplement 147: Second Generation Radiotherapy

5.20.4 Interaction Diagram


[image: image52.emf]Archive PDS or PPMS

C-STORE Monitoring Results)


5.20.4.1 Store RT Patient Positioning Monitoring Result IOD

5.20.4.1.1 Trigger Events

The performing device will finish monitoring, when it receives the signal from the TDD that the beam is switched off. It than stores back the monitoring results.

5.20.4.1.2 Message Semantics

The message semantics are defined by the DICOM Object Storage SOP Classes.

A C-STORE Request shall be sent from the Performing Device to the Archive

5.20.4.1.3 Expected Actions

The archive stores the RT Patient Positioning Monitoring Result IOD.

5.21 RO-DPD-17: Collective UPS Final Update

5.21.1 Scope

In the course of the treatment session, the UPS have been updated by the performing devices with the status and content information based on the result of the UPS execution.

Once the whole treatment session is finished, the TSM propagates the content information for all those UPS, where this was yet done, to the TMS.

5.21.2 Use Case Roles

[image: image53.wmf]Collective UPS Final Update

Treatment Session Manager 

(

TSM

)

Treatment Management System 

(

TMS

)


Actor: Treatment Session Manager ("Performing Device")
Role: Signals using UPS N-SET for each UPS that certain attributes related to the selected work item have changed.
Actor: Treatment Management System

Role: Responds to a UPS N-SET for each UPS and updates attributes in the specified Unified Procedure Step.

5.21.3 Referenced Standards

DICOM 2010 PS 3.4: Annex UUU Unified Procedure Step Service
5.21.4 Interaction Diagram

[image: image54.wmf]TSM

TMS

UPS

-

Pull N

-

SET


5.21.4.1 
UPS Final Update Message

5.21.4.1.1 Trigger Events

The treatment session is finished, when all UPS of  the treatment session are  performed and / or the user indicates to the TSM, that no further activities shall happen and the patient treatment is ended. This causes the TSM to send the final updates of each the UPSs to the TMS.

5.21.4.1.2 Message Semantics

The message semantics for each UPS follows the specification of RO-21.

5.21.4.1.3 Expected Actions

The expected action for each UPS follows the specification of RO-21.

5.22 RO-DPD-18: Collective UPS Completed / Cancelled

5.22.1 Scope

In the course of the treatment session, the UPS have been updated by the performing devices with the status and content information based on the result of the UPS execution.

Once the whole treatment session is finished, the TSM propagates the status for all those UPS, where this was yet done, to the TMS.

5.22.2 Use Case Roles

[image: image55.wmf]Collective UPS Completed 

/ 

Cancelled

Treatment Session Manager 

(

TSM

)

Treatment Management System 

(

TMS

)


Actor: Treatment Session Manager ("Performing Device")
Role: Signals using UPS N-ACTION for each UPS that the selected work item is completed or canceled.
Actor: Treatment Management System

Role: Responds to a UPS N-ACTION for each UPS and sets the specified Unified Procedure Step as completed or canceled.
5.22.3 Referenced Standards

DICOM 2010 PS 3.4: Annex UUU Unified Procedure Step Service
5.22.4 Interaction Diagram

[image: image56.wmf]TSM

TMS

UPS

-

Pull N

-

ACTION


5.22.4.1 Store RT Patient Positioning Monitoring Result IOD

5.22.4.1.1 Trigger Events

When the TSM has finished the final update of all UPS handled in RO-DPD-17, this transaction is invoked.

5.22.4.1.2 Message Semantics

The message semantics for each UPS follows the specification of RO-25.

5.22.4.1.3 Expected Actions

The expected action for each UPS follows the specification of RO-25.

5.23 RO-DPD-19: Indicate Ready for Monitoring

5.23.1 Scope

When monitoring is requested and a beam of the current Treat UPS is selected for delivery, the Monitoring System PPMS indicates once it is ready for monitoring.

5.23.2 Use Case Roles

[image: image57.wmf]TBD

Patient Position Monitoring System 

(

PPMS

)

Treatment Delivery Device 

(

TDD

)


Actor: Patient Positioning Monitoring System ("Performing Device")
Role: Signals using TBD that it is ready for monitoring
Actor: Treatment Delivery Device (TDD)

Role: Uses this signal to get confirmation, that the monitoring system is ready for operation.
5.23.3 Referenced Standards

(TBD)

5.23.4 Interaction Diagram

[image: image58.wmf]PPMS

TDD

(

TBD

)


5.23.4.1 Indicate Ready for Monitoring

5.23.4.1.1 Trigger Events

When the monitoring system has reached a state, that it is ready to start monitoring as soon as the delivery is started, it indicates this state to the TDD.

This indication is conveyed by TBD.

5.23.4.1.2 Message Semantics

The PPMS signals to the TDD its ready state.

5.23.4.1.3 Expected Actions

Once the TDD has received this signal, it will start delivery as soon as its own state does allow this.

5.24 RO-DPD-20: Notify Device to start UPS

5.24.1 Scope

Some actor (in most cases the TSM) notifies a device, that the treatment session process is in a state, that the next UPS can be performed by this device.
5.24.2 Use Case Roles

[image: image59.emf]TBD

Treatment Session Manager (TSM)

Or

Primary Patient Positioning Device (Primary PPD)

Patient Positioning Verification System (PPVS)

Or

Patient Positioning Registration System (PPRS)

Or

Patient Positioning Device (PPD)

Or

Patient Positioning Monitoring System (PPMS)


Actor: Treatment Session Manager or Primary Patient Positioning Device ("Performing Device")
Role: Signals using TBD that it is ready for monitoring
Actor: Patient Positioning Verification System or Patient Positioning Registration or Patient Positioning Monitoring System or Patient Positioning Device 
Role: Uses this signal to get confirmation, that the monitoring system is ready for operation.
5.24.3 Referenced Standards

(TBD)

5.24.4 Interaction Diagram

[image: image60.emf]TSM or PPD PPVS or PPRS or PPD or PPMS

N-EVENT-REPORT

N-EVENT-REPORT Response


5.24.4.1 Indicate Readiness to start UPS
5.24.4.1.1 Trigger Events

The performing device determines, that a UPS is ready to start.

Specifically, this happens for the difference occurrences of this transaction as follows:

5.24.4.1.1.1 UPS to start Patient Position Verification
5.24.4.1.1.2 UPS to start Patient Position Registration

5.24.4.1.1.3 …

((…TBD: collection of triggered events))

5.24.4.1.2 Message Semantics

The performing device sends the subscribed devices the N-EVENT-REPORT as specified in for the Service Class Provider Behavior of DIMSE SERVICE GROUP – UPS Event in Supplement 96.
When the UPS was retrieved by the performing device from another actor (e.g. when the TSM has retrieved the UPS from TMS), this transaction in realized by ((TBD))

When the UPS is created in the course of the treatment session in the performing device, this N-EVENT-REPORT will indicate the creation of the UPS.
5.24.4.1.3 Expected Actions

The subscriber responds with an N-EVENT-REPORT response to indicate the successful reception of the indication for the subscribed UPS changes.

Following this, the subscriber, which is supposed to perform this UPS, will typically continue with the activities like indicated in  the following. Note however, that the following should indicate the typical scenarios, but not part of this transaction any more.
The subscriber system will retrieve the UPS and set it into IN_PROGRESS to get ownership of this UPS. 

Following that, this subscriber will usually start performing the UPS, retrieve associated input data to perform their actions. User interaction with those devices may or may not be needed – this depends on the needs and functionality of those devices. The subscriber may also choose to cancel the UPS in which case the UPS will set into CANCELLED state.
5.25 RO-DPD-21: Create Positioning Acquisition and Positioning Registration UPS

5.25.1 Scope
5.25.2 Use Case Roles

5.25.3 Referenced Standard

5.25.4 Interaction Diagram

5.25.4.1 
Create UPS

5.25.4.1.1 Trigger Events

5.25.4.1.2 Message Semantics

5.25.4.1.3 Expected Actions

5.26 RO-DPD-22: Create Treat UPS and Beam Delivery Instruction for Continuation

5.26.1 Scope

5.26.2 Use Case Roles

5.26.3 Referenced Standard

5.26.4 Interaction Diagram

5.26.4.1 
Create Treat UPS

5.26.4.1.1 Trigger Events

5.26.4.1.2 Message Semantics

5.26.4.1.3 Expected Actions

Appendix General Requirements on using UPS

This appendix contains the shared definition of UPS content for all transactions, which deal with UPS retrieval.

In the this section, the following nomenclature is used;

Performing Device: Represents the actor, which in the specific transactions queries the worklist to pull the UPS

Worklist Server: Represents the actor, which in the specific transaction provides the worklist, the performing device is querying on.

5.27 General Transaction description

The Performing Device uses the C-FIND request of the DICOM Unified Procedure Step – Pull SOP Class to query the worklist from the Worklist Server. The Performing Device performs the SCU role, and the Worklist Server performs the SCP role. Note that the UPS-Pull SOP Class is negotiated as the abstract transfer syntax, and used as the Affected SOP Class in the C-FIND request (see DICOM Supplement 96, Part 4, Section F.X.3.8.1.2.1).

5.28 Matching Keys and Return Keys for Display

In the query to the Worklist Server, the Performing Device (SCU) is required to query for matching on the attributes as shown in Table 3.12-1. All other potential query keys may be optionally supplied as described in DICOM Supplement 96. It is anticipated that Patient’s Name (0010,0010), Patient ID (0010,0020), and Scheduled Station Name Code Sequence (0040,4025) would be optional matching query key attributes that would be commonly supplied.

Table 3.12-1: Matching and Return Keys for UWPS Worklist Query for Positioning

	Attribute Name
	Tag
	Query Keys Matching
	Query Keys Return

	
	
	SCU
	SCP
	SCU
	SCP

	Specific Character Set
	(0008,0005)
	-
	-
	O
	RC

	SOP Class UID
	(0008,1016)
	-
	-
	O
	R

	SOP Instance UID
	(0008,0018)
	-
	-
	O
	R

	Unified Procedure Step State
	(0074,1000)
	R
	R
	R* (Note 1)
	R

	Procedure Step Label
	(0074,1204)
	-
	-
	R+
	R+

	Scheduled Station Name Code Sequence
	(0040,4025)
	
	
	
	

	>Code Value
	(0008,0100)
	O (Note 2)
	R
	R+*
	R

	>Coding Scheme Designator
	(0008,0102)
	O (Note 3)
	R
	R+*
	R

	>Code Meaning 
	(0008,0104)
	-
	-
	R+
	R (Note 4)

	Scheduled Procedure Step Start Date and Time 
	(0040,4005)
	R+ (Note 5)
	R
	R+
	R

	Scheduled Workitem Code Sequence
	(0040,4018)
	-
	-
	R+ (Note 6)
	R

	Scheduled Processing Parameters Sequence
	(0074,1210)
	-
	-
	R+* (Note 7)
	R

	Input Information Sequence
	(0040,4021)
	-
	-
	R+*
	R+ (Note 8)

	Study Instance UID 
	(0020,000D)
	-
	-
	O
	R+ (Note 9)

	Patient’s Name
	(0010,0010)
	O
	R
	R+
	R

	Patient ID
	(0010,0020)
	O
	R
	R+
	R

	All other attributes
	As described in DICOM Supplement 96


Note 1: A Unified Procedure Step State of ‘SCHEDULED’ shall be supplied.

Note 2: Code Value for the Scheduled Station Name shall contain the string used to definitively match the Performing Device instance with its representation on the Worklist Server. It is not necessarily human-readable.

Note 3: Coding Scheme Designator for the Scheduled Station Name is a private coding scheme, and is not used explicitly in this profile.

Note 4: Coding Meaning for the Scheduled Station Name shall contain the human-readable description of the Station Name, and shall be displayed on the Performing Device. Note that this attribute is required by IHE-RO in this profile, but is not required in DICOM Supplement 96.

Note 5: A ‘reasonable’ date time range (such as the rest of the current day) shall be supplied to limit the size of the returned result set. If operating in a mode where the patient is selected on the Worklist Server, the Worklist Server is permitted to over-filter the result set based upon this selection and return just the worklist items for the selected fraction.

Note 6: Scheduled Workitem Code Sequence shall be specified as an empty (null) sequence.

Note 7: Scheduled Processing Parameters Sequence shall be specified as an empty (null) sequence

Note 8: Input Information Sequence shall contain all the input objects that will ultimately be needed to perform the specified procedure step, and no others. This allows the Performing Device to determine whether or not the instances are available prior to starting the procedure, and avoids the need for an additional N-GET on the UPS. If the Performing Device considers that the Input Information Sequence contains inadequate or inconsistent information, then it shall address any such inconsistencies in a safe manner before performing the Requested Procedure.

Note 9: Study Instance UID must be supplied by the Worklist Server (SCP) if performance of the procedure step is expected to create composite SOP Instances as output. The supplied Study Instance shall be used by the SCU in creation of such SOP Instances (see transactions RO-19, RO-20, and RO-21).

In the query to the Worklist Server (SCP), the Performing Device (SCU) is required to query for matching on the attributes shown as “R” or “R+” in the appropriate column in Table 3.12-1. All other potential query keys may be optionally supplied as described in DICOM Supplement 96. It is anticipated that Patient’s Name (0010,0010), Patient ID (0010,0020), and Scheduled Station Name Code Sequence (0040,4025) would be optional matching query key attributes that would be commonly supplied.

In the query to the Worklist Server (SCP), the Performing Device (SCU) is required to supply return keys for display as shown in “Query Keys Return SCU’ column of Table 3.12-1. All other potential return keys for display may be optionally supplied as described in DICOM Supplement 96. The SCU is NOT required to display items marked with an asterisk.

The Worklist Server replies to the query with a set of UPS C-FIND responses containing zero or more scheduled patient positioning worklist items.

5.29 Expected Actions

The Worklist Server retrieves the matching scheduled procedures, and sends the DICOM UWPS Worklist to the requesting Performing Device.

5.30 Receive Scheduled UWPS Worklist Message

This is the message that the Worklist Server sends to the Performing Device as a reply containing DICOM UPS information.

5.31 Trigger Events

The Worklist Server has received a query for a UWPS Worklist.

This profile requires that the Worklist Server has previously received the information needed to effectively respond to queries issued by the Performing Devices that is not generated by the Worklist Server itself. In particular, the Worklist Server must know the SOP Instance UIDs of SOP Instances (such as an RT Plan instance) to be supplied in the Input Information Sequence. It is a precondition of the query transactions that such information has been supplied to the Worklist Server where necessary.

5.32 Message Semantics

C-FIND Responses from the DICOM Unified Procedure Step – Push SOP Class will be used for this message. The Worklist Server will determine the procedure steps required to fulfill the request, based upon the supplied query keys and any additional internal information it may have.

Important: Each matching scheduled UPS shall be returned as a separate UPS C-FIND response item (SOP instance), and each response shall contain exactly one item in the Scheduled Workitem Code Sequence (0040,4018). All procedure steps that need to be performed by the Performing Device shall be returned, i.e. there shall be no additional query for subsequent steps after the first procedure step has been performed.

More than one matching treatment session may be returned by the Worklist Server. For each matching treatment session, a set of Unified Procedure Steps shall be returned, corresponding to the specific procedures that comprise the treatment session. For example, if four matching sessions are found, each composed of four procedures, then the Worklist Server shall return twelve matching Unified Procedure Steps. The specific Scheduled Workitem codes supplied by the Worklist Server shall be consistent with the capabilities of the Performing Device.

<Appendix letter> Appendix_Name

<Detailed cross transaction relationships or mapping details are described in an appendix in Volume 2>

Copyright © 201X: IHE International, Inc.
1
__________________________________________________________________________

1
<Rev. X.X -  201X-MM-DD> 

Copyright © 201X: IHE International, Inc.


_1303917040.vsd
Worklist Query for Positioning


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Verification System (PPVS)


Create UPS for Re-positioning


Treatment Session Manager (TSM)


Patient Positioning Registration System (PPRS)


Store RT Position Modification Instruction


Positioning and Delivery System (PDS) or
Patient Positioning System (PPRS)


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Store Repositioning Results


Archive


Treatment delivery Device (TDD)


Subscribe / Unsubscribe to UPS Treat Event


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Archive


Notify on status change of Treat UPS


Treatment delivery Device (TDD)


Treatment Session Manager (TSM)
or
Patient Position Monitoring System (PPMS)


Retrieve Patient Positioning Device Actuals


Treatment Delivery Device (TDD)


Patient Positioning System (PPVS) or
Patient Positioning System (PPMS)


Worklist Query for Repositioning


Treatment Session Manager (TSM)


Treatment Delivery Device (TDD)


Retrieve RT Position Modification Instruction


Treatment delivery Device (TDD)


Treatment Session Manager (TSM)


Notify on active beam


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Retrieve Positioning Acquisition Results


Archive


Patient Positioning Registration System (PPRS)


Worklist Query for Registration


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Registration System (PPRS)



_1303917515.vsd
Worklist Query for Positioning


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Verification System (PPVS)


Create UPS for Re-positioning


Treatment Session Manager (TSM)


Patient Positioning Registration System (PPRS)


Store RT Position Modification Instruction


Positioning and Delivery System (PDS) or
Patient Positioning System (PPRS)


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Store Repositioning Results


Archive


Treatment delivery Device (TDD)


Subscribe / Unsubscribe to UPS Treat Event


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Archive


Notify on status change of Treat UPS


Treatment delivery Device (TDD)


Treatment Session Manager (TSM)
or
Patient Position Monitoring System (PPMS)


Retrieve Patient Positioning Device Actuals


Treatment Delivery Device (TDD)


Patient Positioning System (PPVS) or
Patient Positioning System (PPMS)


Worklist Query for Repositioning


Treatment Session Manager (TSM)


Treatment Delivery Device (TDD)


Retrieve RT Position Modification Instruction


Treatment Delivery Device (TDD)


Treatment Session Manager (TSM)


Notify on active beam


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Retrieve Positioning Acquisition Results


Archive


Patient Positioning Registration System (PPRS)


Worklist Query for Registration


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Registration System (PPRS)


Worklist Query for Monitoring


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Monitoring System (PPMS)


Retrieve RT Position Modification Instruction


Patient Positioning Device (PPD)


Treatment Delivery Device (TDD)



_1303917805.vsd
Worklist Query for Positioning


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Verification System (PPVS)


Create UPS for Re-positioning


Treatment Session Manager (TSM)


Patient Positioning Registration System (PPRS)


Store RT Position Modification Instruction


Positioning and Delivery System (PDS) or
Patient Positioning System (PPRS)


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Store Repositioning Results


Archive


Treatment Delivery Device (TDD)


Subscribe / Unsubscribe to UPS Treat Event


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Archive


Notify on status change of Treat UPS


Treatment delivery Device (TDD)


Treatment Session Manager (TSM)
or
Patient Position Monitoring System (PPMS)


Retrieve Patient Positioning Device Actuals


Treatment Delivery Device (TDD)


Patient Positioning System (PPVS) or
Patient Positioning System (PPMS)


Worklist Query for Repositioning


Treatment Session Manager (TSM)


Treatment Delivery Device (TDD)


Retrieve RT Position Modification Instruction


Treatment Delivery Device (TDD)


Treatment Session Manager (TSM)


Notify on active beam


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Retrieve Positioning Acquisition Results


Archive


Patient Positioning Registration System (PPRS)


Worklist Query for Registration


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Registration System (PPRS)


Worklist Query for Monitoring


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Monitoring System (PPMS)


Retrieve RT Position Modification Instruction


Patient Positioning Device (PPD)


Treatment Delivery Device (TDD)


Store Repositioning Results


Treatment Delivery Device (TDD)


Patient Positioning Device (TDD)



_1303918121.vsd
TSM

Sequence


PPRS

N-CREATE (UPS)


Sequence


TSM

TDD

Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Archive

TDD

C-STORE (Repositioning Results)


Sequence


Sequence


TDD

PPMS

N-ACTION


Sequence


TDD

TSM or PPMS


N-EVENT_REPORT


Sequence


PDS or PPVS


TMS or TSM


Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Sequence


Sequence


PPVS, PPMS


TDD

Request actual patient position information


Receive actual patient position information


Sequence


TMS or TSM


PDS or PPRS


C-STORE (RT Patient Position Modification Instruction)


Sequence


Sequence


TSM

TDD

C-STORE (RT Patient Position Modification Instruction)


C-MOVE (RT Patient Position Modification Instruction)


Sequence


Sequence


Sequence


TDD

PPMS

N-EVENT-REPORT


Sequence


Sequence


Sequence


Sequence


Archive

PPRS

Retrieve Objects (C-MOVE)


Store Objects (C-STORE)


Sequence


Sequence


PDS or PPRS


TMS or TSM


Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Sequence


Sequence


PDS or PPMS


TMS or TSM


Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Sequence


Sequence


TDD

PPD

C-STORE (RT Patient Position Modification Instruction)


C-MOVE (RT Patient Position Modification Instruction)


Sequence


Sequence


TDD

PPD

C-STORE (Repositioning Results)


Sequence


Sequence


Archive

PDS or PPMS


C-STORE Monitoring Results)



_1326138995.doc


























Transaction-C [C]







Actor D/



Actor E







Actor A /



Actor B







Transaction-B [B]































Internal action1 







Internal action 2 







Transaction 3 [3]







Transaction-A [A]







Transaction 1 [1]











Transaction 2 [2]











Transaction-D [D]







Actor B/ Actor C 







Transaction 2 [2]












_1303918016.vsd
Worklist Query for Positioning


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Verification System (PPVS)


Create UPS for Re-positioning


Treatment Session Manager (TSM)


Patient Positioning Registration System (PPRS)


Store RT Position Modification Instruction


Positioning and Delivery System (PDS) or
Patient Positioning System (PPRS)


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Store Repositioning Results


Archive


Treatment Delivery Device (TDD)


Subscribe / Unsubscribe to UPS Treat Event


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Archive


Notify on status change of Treat UPS


Treatment delivery Device (TDD)


Treatment Session Manager (TSM)
or
Patient Position Monitoring System (PPMS)


Retrieve Patient Positioning Device Actuals


Treatment Delivery Device (TDD)


Patient Positioning System (PPVS) or
Patient Positioning System (PPMS)


Worklist Query for Repositioning


Treatment Session Manager (TSM)


Treatment Delivery Device (TDD)


Retrieve RT Position Modification Instruction


Treatment Delivery Device (TDD)


Treatment Session Manager (TSM)


Notify on active beam


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Retrieve Positioning Acquisition Results


Archive


Patient Positioning Registration System (PPRS)


Worklist Query for Registration


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Registration System (PPRS)


Worklist Query for Monitoring


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Monitoring System (PPMS)


Retrieve RT Position Modification Instruction


Patient Positioning Device (PPD)


Treatment Delivery Device (TDD)


Store Repositioning Results


Treatment Delivery Device (TDD)


Patient Positioning Device (TDD)


Store Monitoring Results


Archive


Positioning and Delivery System (PDS)
or
Patient Positioning Monitoring System (PPMS)



_1303917704.vsd
TSM

Sequence


PPRS

N-CREATE (UPS)


Sequence


TSM

TDD

Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Archive

TDD

C-STORE (Repositioning Results)


Sequence


Sequence


TDD

PPMS

N-ACTION


Sequence


TDD

TSM or PPMS


N-EVENT_REPORT


Sequence


PDS or PPVS


TMS or TSM


Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Sequence


Sequence


PPVS, PPMS


TDD

Request actual patient position information


Receive actual patient position information


Sequence


TMS or TSM


PDS or PPRS


C-STORE (RT Patient Position Modification Instruction)


Sequence


Sequence


TSM

TDD

C-STORE (RT Patient Position Modification Instruction)


C-MOVE (RT Patient Position Modification Instruction)


Sequence


Sequence


Sequence


TDD

PPMS

N-EVENT-REPORT


Sequence


Sequence


Sequence


Sequence


Archive

PPRS

Retrieve Objects (C-MOVE)


Store Objects (C-STORE)


Sequence


Sequence


PDS or PPRS


TMS or TSM


Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Sequence


Sequence


PDS or PPMS


TMS or TSM


Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Sequence


Sequence


TDD

PPD

C-STORE (RT Patient Position Modification Instruction)


C-MOVE (RT Patient Position Modification Instruction)


Sequence


Sequence


TDD

PPD

C-STORE (Repositioning Results)


Sequence



_1303917780.vsd
Worklist Query for Positioning


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Verification System (PPVS)


Create UPS for Re-positioning


Treatment Session Manager (TSM)


Patient Positioning Registration System (PPRS)


Store RT Position Modification Instruction


Positioning and Delivery System (PDS) or
Patient Positioning System (PPRS)


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Store Repositioning Results


Archive


Treatment delivery Device (TDD)


Subscribe / Unsubscribe to UPS Treat Event


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Archive


Notify on status change of Treat UPS


Treatment delivery Device (TDD)


Treatment Session Manager (TSM)
or
Patient Position Monitoring System (PPMS)


Retrieve Patient Positioning Device Actuals


Treatment Delivery Device (TDD)


Patient Positioning System (PPVS) or
Patient Positioning System (PPMS)


Worklist Query for Repositioning


Treatment Session Manager (TSM)


Treatment Delivery Device (TDD)


Retrieve RT Position Modification Instruction


Treatment Delivery Device (TDD)


Treatment Session Manager (TSM)


Notify on active beam


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Retrieve Positioning Acquisition Results


Archive


Patient Positioning Registration System (PPRS)


Worklist Query for Registration


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Registration System (PPRS)


Worklist Query for Monitoring


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Monitoring System (PPMS)


Retrieve RT Position Modification Instruction


Patient Positioning Device (PPD)


Treatment Delivery Device (TDD)


Store Repositioning Results


Treatment Delivery Device (TDD)


Patient Positioning Device (TDD)



_1303917616.vsd
TSM

Sequence


PPRS

N-CREATE (UPS)


Sequence


TSM

TDD

Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Archive

TDD

C-STORE (Repositioning Results)


Sequence


Sequence


TDD

PPMS

N-ACTION


Sequence


TDD

TSM or PPMS


N-EVENT_REPORT


Sequence


PDS or PPVS


TMS or TSM


Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Sequence


Sequence


PPVS, PPMS


TDD

Request actual patient position information


Receive actual patient position information


Sequence


TMS or TSM


PDS or PPRS


C-STORE (RT Patient Position Modification Instruction)


Sequence


Sequence


TSM

TDD

C-STORE (RT Patient Position Modification Instruction)


C-MOVE (RT Patient Position Modification Instruction)


Sequence


Sequence


Sequence


TDD

PPMS

N-EVENT-REPORT


Sequence


Sequence


Sequence


Sequence


Archive

PPRS

Retrieve Objects (C-MOVE)


Store Objects (C-STORE)


Sequence


Sequence


PDS or PPRS


TMS or TSM


Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Sequence


Sequence


PDS or PPMS


TMS or TSM


Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Sequence


Sequence


TDD

PPD

C-STORE (RT Patient Position Modification Instruction)


C-MOVE (RT Patient Position Modification Instruction)


Sequence


Sequence



_1303917311.vsd
Worklist Query for Positioning


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Verification System (PPVS)


Create UPS for Re-positioning


Treatment Session Manager (TSM)


Patient Positioning Registration System (PPRS)


Store RT Position Modification Instruction


Positioning and Delivery System (PDS) or
Patient Positioning System (PPRS)


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Store Repositioning Results


Archive


Treatment delivery Device (TDD)


Subscribe / Unsubscribe to UPS Treat Event


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Archive


Notify on status change of Treat UPS


Treatment delivery Device (TDD)


Treatment Session Manager (TSM)
or
Patient Position Monitoring System (PPMS)


Retrieve Patient Positioning Device Actuals


Treatment Delivery Device (TDD)


Patient Positioning System (PPVS) or
Patient Positioning System (PPMS)


Worklist Query for Repositioning


Treatment Session Manager (TSM)


Treatment Delivery Device (TDD)


Retrieve RT Position Modification Instruction


Treatment delivery Device (TDD)


Treatment Session Manager (TSM)


Notify on active beam


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Retrieve Positioning Acquisition Results


Archive


Patient Positioning Registration System (PPRS)


Worklist Query for Registration


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Registration System (PPRS)


Worklist Query for Monitoring


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Monitoring System (PPMS)



_1303917496.vsd
Worklist Query for Positioning


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Verification System (PPVS)


Create UPS for Re-positioning


Treatment Session Manager (TSM)


Patient Positioning Registration System (PPRS)


Store RT Position Modification Instruction


Positioning and Delivery System (PDS) or
Patient Positioning System (PPRS)


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Store Repositioning Results


Archive


Treatment delivery Device (TDD)


Subscribe / Unsubscribe to UPS Treat Event


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Archive


Notify on status change of Treat UPS


Treatment delivery Device (TDD)


Treatment Session Manager (TSM)
or
Patient Position Monitoring System (PPMS)


Retrieve Patient Positioning Device Actuals


Treatment Delivery Device (TDD)


Patient Positioning System (PPVS) or
Patient Positioning System (PPMS)


Worklist Query for Repositioning


Treatment Session Manager (TSM)


Treatment Delivery Device (TDD)


Retrieve RT Position Modification Instruction


Treatment Delivery Device (TDD)


Treatment Session Manager (TSM)


Notify on active beam


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Retrieve Positioning Acquisition Results


Archive


Patient Positioning Registration System (PPRS)


Worklist Query for Registration


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Registration System (PPRS)


Worklist Query for Monitoring


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Monitoring System (PPMS)


Retrieve RT Position Modification Instruction


Patient Positioning Device (PPD)


Treatment Delivery Device (TDD)



_1303917102.vsd
TSM

Sequence


PPRS

N-CREATE (UPS)


Sequence


TSM

TDD

Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Archive

TDD

C-STORE (Repositioning Results)


Sequence


Sequence


TDD

PPMS

N-ACTION


Sequence


TDD

TSM or PPMS


N-EVENT_REPORT


Sequence


PDS or PPVS


TMS or TSM


Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Sequence


Sequence


PPVS, PPMS


TDD

Request actual patient position information


Receive actual patient position information


Sequence


TMS or TSM


PDS or PPRS


C-STORE (RT Patient Position Modification Instruction)


Sequence


Sequence


TSM

TDD

C-STORE (RT Patient Position Modification Instruction)


C-MOVE (RT Patient Position Modification Instruction)


Sequence


Sequence


Sequence


TDD

PPMS

N-EVENT-REPORT


Sequence


Sequence


Sequence


Sequence


Archive

PPRS

Retrieve Objects (C-MOVE)


Store Objects (C-STORE)


Sequence


Sequence


PDS or PPRS


TMS or TSM


Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Sequence


Sequence


Sequence



_1303917167.vsd
TSM

Sequence


PPRS

N-CREATE (UPS)


Sequence


TSM

TDD

Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Archive

TDD

C-STORE (Repositioning Results)


Sequence


Sequence


TDD

PPMS

N-ACTION


Sequence


TDD

TSM or PPMS


N-EVENT_REPORT


Sequence


PDS or PPVS


TMS or TSM


Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Sequence


Sequence


PPVS, PPMS


TDD

Request actual patient position information


Receive actual patient position information


Sequence


TMS or TSM


PDS or PPRS


C-STORE (RT Patient Position Modification Instruction)


Sequence


Sequence


TSM

TDD

C-STORE (RT Patient Position Modification Instruction)


C-MOVE (RT Patient Position Modification Instruction)


Sequence


Sequence


Sequence


TDD

PPMS

N-EVENT-REPORT


Sequence


Sequence


Sequence


Sequence


Archive

PPRS

Retrieve Objects (C-MOVE)


Store Objects (C-STORE)


Sequence


Sequence


PDS or PPRS


TMS or TSM


Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Sequence


Sequence


PDS or PPMS


TMS or TSM


Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Sequence


Sequence



_1303915345.vsd
Worklist Query for Positioning


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning System (PPVS)


Create UPS for Re-positioning


Treatment Session Manager (TSM)


Patient Positioning Registration System (PPRS)


Store RT Position Modification Instruction


Treatment Session Manager (TSM)


Patient Positioning System (PPVS)


Store Repositioning Results


Archive


Treatment delivery Device (TDD)


Subscribe / Unsubscribe to UPS Treat Event


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Archive


Notify on status change of Treat UPS


Treatment delivery Device (TDD)


Treatment Session Manager (TSM)
or
Patient Position Monitoring System (PPMS)


Retrieve Patient Positioning Device Actuals


Treatment Delivery Device (TDD)


Patient Positioning System (PPVS) or
Patient Positioning System (PPMS)


Worklist Query for Repositioning


Treatment Session Manager (TSM)


Treatment Delivery Device (TDD)


Retrieve RT Position Modification Instruction


Treatment delivery Device (TDD)


Treatment Session Manager (TSM)


Notify on active beam


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Retrieve Positioning Acquisition Results


Archive


Patient Positioning Registration System (PPRS)



_1303915750.vsd
TSM

Sequence


PPRS

N-CREATE (UPS)


Sequence


TSM

TDD

Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


TMS

TDD

C-STORE (Repositioning Results)


Sequence


Sequence


TDD

PPMS

N-ACTION


Sequence


TDD

TSM or PPMS


N-EVENT_REPORT


Sequence


PDS or PPVS


TMS or TSM


Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Sequence


Sequence


PPVS, PPMS


TDD

Request actual patient position information


Receive actual patient position information


Sequence


TMS or TSM


PDS or PPRS

C-STORE (RT Patient Position Modification Instruction)


Sequence


Sequence


TSM

TDD

C-STORE (RT Patient Position Modification Instruction)


C-MOVE (RT Patient Position Modification Instruction)


Sequence


Sequence


Sequence


TDD

PPMS

N-EVENT-REPORT


Sequence


Sequence


Sequence


Sequence


Archive

PPRS

Retrieve Objects (C-MOVE)


Store Objects (C-STORE)



_1303915953.vsd
TSM

Sequence


PPRS

N-CREATE (UPS)


Sequence


TSM

TDD

Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Archive

TDD

C-STORE (Repositioning Results)


Sequence


Sequence


TDD

PPMS

N-ACTION


Sequence


TDD

TSM or PPMS


N-EVENT_REPORT


Sequence


PDS or PPVS


TMS or TSM


Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Sequence


Sequence


PPVS, PPMS


TDD

Request actual patient position information


Receive actual patient position information


Sequence


TMS or TSM


PDS or PPRS

C-STORE (RT Patient Position Modification Instruction)


Sequence


Sequence


TSM

TDD

C-STORE (RT Patient Position Modification Instruction)


C-MOVE (RT Patient Position Modification Instruction)


Sequence


Sequence


Sequence


TDD

PPMS

N-EVENT-REPORT


Sequence


Sequence


Sequence


Sequence


Archive

PPRS

Retrieve Objects (C-MOVE)


Store Objects (C-STORE)



_1303917017.vsd
Worklist Query for Positioning


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning System (PPVS)


Create UPS for Re-positioning


Treatment Session Manager (TSM)


Patient Positioning Registration System (PPRS)


Store RT Position Modification Instruction


Positioning and Delivery System (PDS) or
Patient Positioning System (PPRS)


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Store Repositioning Results


Archive


Treatment delivery Device (TDD)


Subscribe / Unsubscribe to UPS Treat Event


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Archive


Notify on status change of Treat UPS


Treatment delivery Device (TDD)


Treatment Session Manager (TSM)
or
Patient Position Monitoring System (PPMS)


Retrieve Patient Positioning Device Actuals


Treatment Delivery Device (TDD)


Patient Positioning System (PPVS) or
Patient Positioning System (PPMS)


Worklist Query for Repositioning


Treatment Session Manager (TSM)


Treatment Delivery Device (TDD)


Retrieve RT Position Modification Instruction


Treatment delivery Device (TDD)


Treatment Session Manager (TSM)


Notify on active beam


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Retrieve Positioning Acquisition Results


Archive


Patient Positioning Registration System (PPRS)


Worklist Query for Registration


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning Registration System (PPRS)



_1303915697.vsd
Worklist Query for Positioning


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning System (PPVS)


Create UPS for Re-positioning


Treatment Session Manager (TSM)


Patient Positioning Registration System (PPRS)


Store RT Position Modification Instruction


Positioning and Delivery System (PDS) or
Patient Positioning System (PPRS)


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Store Repositioning Results


Archive


Treatment delivery Device (TDD)


Subscribe / Unsubscribe to UPS Treat Event


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Archive


Notify on status change of Treat UPS


Treatment delivery Device (TDD)


Treatment Session Manager (TSM)
or
Patient Position Monitoring System (PPMS)


Retrieve Patient Positioning Device Actuals


Treatment Delivery Device (TDD)


Patient Positioning System (PPVS) or
Patient Positioning System (PPMS)


Worklist Query for Repositioning


Treatment Session Manager (TSM)


Treatment Delivery Device (TDD)


Retrieve RT Position Modification Instruction


Treatment delivery Device (TDD)


Treatment Session Manager (TSM)


Notify on active beam


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Retrieve Positioning Acquisition Results


Archive


Patient Positioning Registration System (PPRS)



_1303913914.vsd
TSM

Sequence


PPVS

N-CREATE (UPS)


Sequence


TSM

TDD

Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


TMS

TDD

C-STORE (Repositioning Results)


Sequence


Sequence


TDD

PPMS

N-ACTION


Sequence


TDD

TSM or PPMS


N-EVENT_REPORT


Sequence


PDS or PPVS


TMS or TSM


Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Sequence


Sequence


PPVS, PPMS


TDD

Request actual patient position information


Receive actual patient position information


Sequence


TSM or TMS


PPVS

C-STORE (RT Patient Position Modification Instruction)


Sequence


Sequence


TSM

TDD

C-STORE (RT Patient Position Modification Instruction)


C-MOVE (RT Patient Position Modification Instruction)


Sequence


Sequence


Sequence


TDD

PPMS

N-EVENT-REPORT


Sequence


Sequence


Sequence


Sequence


Archive

PPRS

Retrieve Objects (C-MOVE)


Store Objects (C-STORE)



_1303915211.vsd
TSM

Sequence


PPRS

N-CREATE (UPS)


Sequence


TSM

TDD

Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


TMS

TDD

C-STORE (Repositioning Results)


Sequence


Sequence


TDD

PPMS

N-ACTION


Sequence


TDD

TSM or PPMS


N-EVENT_REPORT


Sequence


PDS or PPVS


TMS or TSM


Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Sequence


Sequence


PPVS, PPMS


TDD

Request actual patient position information


Receive actual patient position information


Sequence


TSM or TMS


PPVS

C-STORE (RT Patient Position Modification Instruction)


Sequence


Sequence


TSM

TDD

C-STORE (RT Patient Position Modification Instruction)


C-MOVE (RT Patient Position Modification Instruction)


Sequence


Sequence


Sequence


TDD

PPMS

N-EVENT-REPORT


Sequence


Sequence


Sequence


Sequence


Archive

PPRS

Retrieve Objects (C-MOVE)


Store Objects (C-STORE)



_1303893786.vsd
Worklist Query for Positioning


Treatment Management System (TMS)
or
Treatment Session Manager (TSM)


Positioning and Delivery System (PDS)
or
Patient Positioning System (PPVS) or
Patient Positioning System (PPMS)


Create UPS for Re-positioning


Treatment Session Manager (TSM)


Patient Positioning System (PPVS)


Store RT Position Modification Instruction


Treatment Session Manager (TSM)


Patient Positioning System (PPVS)


Store Repositioning Results


Archive


Treatment delivery Device (TDD)


Subscribe / Unsubscribe to UPS Treat Event


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Archive


Notify on status change of Treat UPS


Treatment delivery Device (TDD)


Treatment Session Manager (TSM)
or
Patient Position Monitoring System (PPMS)


Retrieve Patient Positioning Device Actuals


Treatment Delivery Device (TDD)


Patient Positioning System (PPVS) or
Patient Positioning System (PPMS)


Worklist Query for Repositioning


Treatment Session Manager (TSM)


Treatment Delivery Device (TDD)


Retrieve RT Position Modification Instruction


Treatment delivery Device (TDD)


Treatment Session Manager (TSM)


Notify on active beam


Treatment delivery Device (TDD)


Patient Position Monitoring System (PPMS)


Retrieve Positioning Acquisition Results


Archive


Patient Positioning Registration System (PPRS)



_1303894321.vsd
TSM

Sequence


PPVS

N-CREATE (UPS)


Sequence


TSM

TDD

Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


TMS

TDD

C-STORE (Repositioning Results)


Sequence


Sequence


TDD

PPMS

N-ACTION


Sequence


TDD

TSM or PPMS


N-EVENT_REPORT


Sequence


PPVS, PPMS or PDS


TMS or TSM


Query Scheduled UPS Worklist


Receive Scheduled UPS Worklist


Sequence


Sequence


Sequence


PPVS, PPMS


TDD

Request actual patient position information


Receive actual patient position information


Sequence


TSM or TMS


PPVS

C-STORE (RT Patient Position Modification Instruction)


Sequence


Sequence


TSM

TDD

C-STORE (RT Patient Position Modification Instruction)


C-MOVE (RT Patient Position Modification Instruction)


Sequence


Sequence


Sequence


TDD

PPMS

N-EVENT-REPORT


Sequence


Sequence


Sequence


Sequence


Archive

PPRS

Retrieve Objects (C-MOVE)


Store Objects (C-STORE)



_1112784916.doc














































( Transaction _2 [2]







( Transaction _2 [2]







Transaction_1 [1] (







































Transaction_1 [1] (







( Transaction_3 [3]



( Transaction_4 [4]











Actor A







Actor F



















Actor E







Actor C







Actor D



































































































































































































































































































































































Actor B




















































